# Empanelment of Equipment’s

Empanelment

Project Management Unit,

National Hydrology Project

Central Water Commission (CWC)

Country: India

Project: National Hydrology Project

INVITATION FOR PROPOSALS (IFP) FOR EMPANELMENT OF

HYDRO-METEOROLOGICAL AND WATER QUALITY EQUIPMENT EMPANELMENT

| IFP REFERENCE | : | NHP/2016-RDD/01 |
| --- | --- | --- |
| DATE OF COMMENCEMENT OF DOWNLOADING OF DOCUMENT FOR SETTING-UP EMPANELMENT AGREEMENT | : | 28.09.2016 at 10:00 AM |
| LAST DATE FOR DOWNLOADING OF DOCUMENT FOR SETTING-UP EMPANELMENT AGREEMENT | : | 10.11.2016 at 18:00 Hrs. |
| LAST DATE AND TIME FOR UPLOADING/ RECEIPT ON PORTAL | : | 11.11.2016 upto 15:00 Hrs. |
| DATE AND TIME OF OPENING OF PROPOSALS | : | 11.11.2016 at 16:00 Hrs. |
| PLACE OF OPENING OF PROPOSALS | : | Project Management Unit, National Hydrology Project, Director, River Data Directorate, Central Water Water Commission, West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605 |
| ADDRESS FOR COMMUNICATION | : : | Director, River Data Directorate, Central Water Commission, West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605 E-mail:- rdcdte-cwc@nic.in, Fax No+91-11-26181267, Telephone + 91-11-26100285/26108075) |

# Invitation for Proposals

Invitation for Proposals (IFP)

India

National Hydrology Project.

IFP title: Empanelment Agreement for Hydro-Meteorological and Water Quality Equipment

IFP number HO: CPMU:NHP/2016-RDD/01

- The Government of India [GOI] has applied for a Loan from the International Bank for Reconstruction and Development toward the cost of National Hydrology Project (NHP) and intends to apply part of the proceeds of this Loan to payments for the Procurements to be issued separately Empanelment based on the empanelment of equipment’s  to be set up through this IFP.

- NHP is countrywide project, with 49 Implementing agencies (IAs) including eight central agencies, 39 state-UT level agencies and two river basin organizations (RBO)

- The Project Management Unit, NHP, CWC now invites sealed Proposals from eligible manufacturers and suppliers who wish their equipment/s to be on the Empaneled Agreement (EA) for supply of Hydro-meteorological and Water Quality equipment.  The manufacturers and Authorized Representatives on the behalf of manufacturer (herein called as ‘Proposer’ in this document) may submit Proposal for one or more items or multiple models of same item. However, evaluation and finalization of the Empanelment Agreements will be done individually for each item. Prices of items are not to be submitted at this stage.

- During the First stage Empanelment agreement, only the makes and models of respective equipment would be evaluated against desired specifications and would be included in the empanelment. After the first stage, all the equipment’s and models which meet the Empanelment evaluation and qualification criteria specified in this document would be included on the empanelled list. The List of all the equipment, makes and models empanelled would be uploaded on the project website.

- During the second stage, bids / quotations would be invited from bidders (integrators) following the Bank guidelines for supply, installation, operation and maintenance of the Hydro-meteorological equipment. The Bidders would be informed about the list of empanelled equipment’s and requested to quote from the list of equipment empanelled which are meeting the purchaser’s technical requirements for each of the equipment’s.

- Contracts/ Purchase Orders will be placed to the Bidder (integrators) that offers lowest price and comply with delivery requirements (Offering only those make/model of the equipment already accepted in first stage Empanelment). During the second stage price competition to be carried out by End Implementing Agencies (EIAs) of the Project. The Empanelment will impose no obligation on the Project Management Unit, National Hydrology Project (NHP), CWC or the EIAs to purchase the estimated or any quantity from the Proposers Empanelment.

- The Empanelled Empanelment shall be valid for a period of 3 year. The validity may be extended for a further period (maximum 2 years) Empanelment. This being open Empanelment new Proposers may submit their Proposals (in the format given in original document for Empanelment for their equipment) any time after initial empanelment. Such offers received after initial Empanelment will be opened at least once in a year (the date of empanelment would be advertise and also be intimated on the website of NHP) and will be evaluated on the basis of qualification and evaluation criteria specified in the original document for setting up the Empanelment.

- Competition will be conducted through the Empanelment Agreement Procedures and is open to all Proposers from eligible source countries as defined in the World Bank’s Guidelines: Procurement under IBRD Loans and IDA Credits.

- Interested eligible Proposers may obtain further information from Project Management Unit, National Hydrology Project (NHP) and inspect the documents for setting up Empanelment agreement at the Committee Room, Central Water Commission, 3rd Floor, Sewa Bhawan, R.K.Puram, New Delhi. A pre-proposal meeting will be held on 17.10.2016 at 11:00 hrs. at the address given  to provide additional clarifications. Non-attendance at the pre-proposal meeting will not be a cause for disqualification of a Proposer. Interested Proposers should depute their staff only to attend the pre- proposal meeting.

- Documents for setting up Empanelment may also be viewed and downloaded from the website www.mowr.gov.in/ www.cwc.gov.in/ www.eprocure.gov.in/ www.indiawrm.org /www.tcil-india-electronictender.com free of cost. In such cases the Proposer would be solely responsible for ensuring that any subsequent addenda issued thereafter and available in website is also downloaded / incorporated in the document while preparing and submitting Proposals. However, in order to be able to participate in the tender it is mandatory to download official copy of tenders from https://www.tcil-india-electronictender.com.

- The bidder has to submit the proposal participation fee of  INR 15000/- in the form of A/C payee Demand Draft (non - refundable) drawn on any scheduled bank in favour of the Deputy Director, River Data Directorate, Central Water Commission, New Delhi payable at New Delhi in a separate envelope to Director, River Data Directorate, Central Water Commission, West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605 (E-mail:- rdcdte-cwc@nic.in, Fax No+91-11-26181267, Telephone + 91-11-26100285/26108075) before the opening date and time of Empanelment document.

- Proposals must be submitted online only at TCIL website URL https://www.tcil-india-electronictender.com before 11.11.2016 upto 15:00 hrs. Tenderer/Contractor is advised to follow the instructions provided in the 'Instructions to the Contractor/Tenderer for the       e-submission of the bids online through TCIL portal. All Proposals must be accompanied by a proposal security as specified in the documents for setting up Empanelment. The proposal security must be reach in a separate envelope to Director, River Data Directorate, Central Water Commission, West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605 (E-mail:- rdcdte-cwc@nic.in, Fax No +91-11-26181267, Telephone + 91-11-26100285/26108075) before the opening date and time of Empanelment document. Late Proposals will be rejected, and they can submit their proposals in the next round of empanelment.  Proposals will be e-opened.

- Intending tenderers are advised to visit again TCIL website URL https://www.tcil-india-electronictender.com, and Ministry of Water Resources, RD &GR website www.mowr.gov.in, CWC website www.cwc.nic.in and CPPP website https://eprocure.gov.in  at least 1 day prior to closing date of submission of tender for any corrigendum /amendment.

- In the event of the date being declared as a closed holiday for purchaser’s office, the due date for submission of Proposals and opening will be carried out on the date and time as notified by the purchaser. The necessary corrigendum shall be issued in this regard by the purchaser.

- SPECIAL INSTRUCTIONS TO BIDDERS BY TCIL FOR E - TENDERING

The Special Instructions (for e-Tendering) supplement ‘Instruction to Bidders’, as given in these Tender Documents. Submission of Online Bids is mandatory for this Tender.

E-Tendering is a new methodology for conducting Public Procurement in a transparent and secured manner. Now, the Government of India has made e-tendering mandatory. Suppliers/ Vendors will be the biggest beneficiaries of this new system of procurement. For conducting electronic tendering, Central Water Commission has decided to use the portal https://www.tcil-indiaelectronictender.com through TCIL, a Government of India Undertaking. This portal is based on the world’s most ‘secure’ and ‘user friendly’ software from Electronic Tender®. A portal built using Electronic Tender’s Software is also referred to as Electronic Tender System® (ETS).

Benefits to tenderers are outlined on the Home-page of the portal.

Instructions

Tender Bidding Methodology:

Single Stage Envelope

Broad Outline of Activities from Bidder’s Perspective:

Procure a Digital Signing Certificate (DSC)

Register on Electronic Tendering System® (ETS)

Create Marketing Authorities (MAs),

Users and assign roles on ETS

View Notice Inviting Tender (NIT) on ETS

For this tender -- Assign Tender Search Code (TSC) to a MA

Download Official Copy of Tender Documents from ETS

Clarification to Tender Documents on ETS

Query to Central Water Commission (Optional)

View response to queries posted by Central Water Commission

Bid-Submission on ETS

Attend Public Online Tender Opening Event (TOE) on ETS

– Opening of relevant Bid-Part

Post-TOE Clarification on ETS (Optional)

– Respond to Central Water Commission Post-TOE queries

Attend Public Online Tender Opening Event (TOE) on ETS

Opening of relevant part (i.e. Financial-Part)

(Only for Technical Responsive Bidders)

Participate in e-Reverse Auction on ETS

For participating in this tender online, the following instructions are to be read carefully.

These instructions are supplemented with more detailed guidelines on the relevant screens of the ETS.

Digital Certificates

For integrity of data and authenticity/non-repudiation of electronic records, and to be compliant with IT Act 2000, it is necessary for each user to have a Digital Certificate (DC). Also referred to as Digital Signature Certificate (DSC), of Class 2 or above, issued by a Certifying Authority (CA) licensed by Controller of Certifying Authorities (CCA) [refer http://www.cca.gov.in].

Registration

To use the Electronic Tender® portal https://www.tcil-india-electronictender.com, vendors need to register on the portal. Registration of each organization is to be done by one of its senior persons who will be the main person coordinating for the etendering activities. In ETS terminology, this person will be referred to as the Super User (SU) of that organization. For further details, please visit the website/portal, and click on the ‘Supplier Organization’ link under ‘Registration’ (on the Home Page), and follow further instructions as given on the site. Pay Annual Registration Fee as applicable.

After successful submission of Registration details and Annual Registration Fee, please contact      TCIL/ETS Helpdesk (as given below), to get your registration accepted/activated.

Important Note: To minimize teething problems during the use of ETS (including the  Registration process), it is recommended that the user should peruse the instructions given under ‘ETS User-Guidance Center’ located on ETS Home Page, including instructions for timely registration on ETS. The instructions relating to ‘Essential Computer Security Settings for Use of ETS’ and ‘Important Functionality Checks’ should be especially taken into cognizance.

Please note that even after acceptance of your registration by the Service Provider, to respond to a tender you will also require time to complete activities related to your organization, such as creation of users, assigning roles to them, etc.

TCIL/ ETS Helpdesk Telephone/ Mobile  Customer Support (0930 hrs to 1800 hrs, Monday to Friday except on gazetted holidays): +91-1126202699  (Multiple lines)  Emergency Support Mobile Numbers: +919868393775, 9868393717, 9868393792  E-mail ID  ets_support@tcil-india.com

Some Bidding related Information for this Tender (Sealed Bid)

The entire bid-submission would be online on ETS (unless specified for Offline Submissions). Broad outline of submissions are as follows:  Submission of Bid-Parts/ Envelopes Single-Part, Submission of information pertaining Bid Security/ Earnest Money Deposit (EMD),  Submission of digitally signed copy of Tender Documents/ Addendum, Submission of General Terms and Conditions (with/ without deviations) and  Submission of Special Terms and Conditions (with/ without deviations).

Special Note on Security and Transparency of Bids

Security related functionality has been rigorously implemented in ETS in a multidimensional manner. Starting with 'Acceptance of Registration by the Service Provider', provision for security has been made at various stages in Electronic Tender's software. Specifically for Bid Submission, some security related aspects are outlined below:  As part of the Electronic Encrypter™ functionality, the contents of both the ‘Electronic Forms’ and the ‘Main-Bid’ are securely encrypted using a Pass-Phrase created by the Bidder himself. Unlike a ‘password’, a Pass-Phrase can be a multi-word sentence with spaces between words (e.g. I love this World). A Pass-Phrase is easier to remember, and more difficult to break. It is recommended that a separate Pass-Phrase be created for each Bid-Part. This method of bid-encryption does not have the security and data-integrity related vulnerabilities which are inherent in e-tendering systems which use Public-Key of the specified officer of a Buyer organization for bid-encryption. Bid-encryption in ETS is such that the Bids cannot be decrypted before the Public Online Tender Opening Event (TOE), even if there is connivance between the concerned tender-opening officers of the Buyer organization and the personnel of e-tendering service provider.

CAUTION: All bidders must fill Electronic Forms™ for each bid-part sincerely and carefully, and avoid any discrepancy between information given in the Electronic Forms™ and the corresponding Main-Bid. For transparency, the information submitted by a bidder in the Electronic Forms™ is made available to other bidders during the Online Public TOE. If it is found during the Online Public TOE that a bidder has not filled in the complete information in the Electronic Forms™, the TOE officer may make available for downloading the corresponding Main-Bid of that bidder at the risk of the bidder. If variation is noted between the information contained in the Electronic Forms™ and the ‘Main-Bid’, the contents of the Electronic Forms™ shall prevail. Alternatively, the Buyer organization reserves the right to consider the higher of the two pieces of information (e.g. the higher price) for the purpose of short-listing, and the lower of the two pieces of information (e.g. the lower price) for the purpose of payment in case that bidder is an awardee in that tender.

Typically, ‘Pass-Phrase’ of the Bid-Part to be opened during a particular Public Online Tender Opening Event (TOE) is furnished online by each bidder during the TOE itself, when demanded by the concerned Tender Opening Officer.

(Optional Text in EBI, depending upon the decision of the Buyer organization):  Additionally, the bidder shall make sure that the Pass-Phrase to decrypt the relevant Bid-Part is submitted to Buyer Organization Name in a sealed envelope before the start date and time of the Tender Opening Event (TOE).

OR

Additionally, the bidder shall make sure that the Pass-Phrase to decrypt the relevant Bid-Part is submitted into the ‘Time Locked Electronic Key Box (EKB)’ after the corresponding deadline of Bid Submission, and before the commencement of the Online TOE. The process of submission of this Pass-Phrase in the ‘Time Locked Electronic Key Box’ is done in a secure manner by first encrypting this Pass-Phrase with the designated keys provided by the Buyer organization.  There is an additional protection with SSL Encryption during transit from the client end computer of a Supplier organization to the e-tendering server/ portal.  Public Online Tender Opening Event (TOE) ETS offers a unique facility for ‘Public Online Tender Opening Event (TOE)’. Tender Opening Officers, as well as, authorized representatives of bidders can simultaneously attend the Public Online Tender Opening Event (TOE) from the comfort of their offices. Alternatively, one/ two duly authorized representative(s) of bidders (i.e. Supplier organization) are requested to carry a Laptop with Wireless Internet Connectivity, if they wish to come to Buyer Organization Name office for the Public Online TOE.  Every legal requirement for a transparent and secure ‘Public Online Tender Opening Event (TOE)’, including digital counter-signing of each opened bid by the authorized TOE Event  (TOE)’, including digital counter-signing of each opened bid by the authorized TOE officer(s)in the simultaneous online presence of the participating bidders’ representatives, has been implemented on ETS.

As soon as a Bid is decrypted with the corresponding ‘Pass-Phrase’ as submitted online by the bidder himself (during the TOE itself), salient points of the Bids (as identified by the Buyer organization) are simultaneously made available for downloading by all participating bidders. The tedium of taking notes during a manual ‘Tender Opening Event’ is therefore replaced with this superior and convenient form of ‘Public Online Tender Opening Event (TOE)’.  ETS has a unique facility of ‘Online Comparison Chart’ which is dynamically updated as each online bid is opened. The format of the chart is based on inputs provided by the Buyer for each Bid-Part of a tender. The information in the Comparison Chart is based on the data submitted by the Bidders. A detailed Technical and/ or Financial Comparison Chart enhances Transparency. Detailed instructions are given on relevant screens.  ETS has a unique facility of a detailed report titled ‘Minutes of Online Tender Opening Event (TOE)’ covering all important activities of ‘Online Tender Opening Event (TOE)’. This is available to all participating bidders for ‘Viewing/ Downloading’.  There are many more facilities and features on ETS. For a particular tender, the screens viewed by a Supplier will depend upon the options selected by the concerned Buyer.

Other Instructions

For further instructions, the vendor should visit the home-page of the portal https://www.tcil-india-electronictender.com, and go to the User-Guidance Center  The help information provided through ‘ETS User-Guidance Center’ is available in three categories – Users intending to Register / First-Time Users, Logged-in users of Buyer organizations, and Logged-in users of Supplier organizations. Various links (including links for User Manuals) are provided under each of the three categories.  Important Note: It is strongly recommended that all authorized users of Supplier organizations should thoroughly peruse the information provided under the relevant links, and take appropriate action. This will prevent hiccups, and minimize teething problems during the use of ETS.

SIX CRITICAL DO’S AND DON’TS FOR BIDDERS

Specifically for Supplier organizations, the following 'SIX KEY INSTRUCTIONS for BIDDERS' must be assiduously adhered to:

1. Obtain individual Digital Signing Certificate (DSC or DC) well in advance of your first tender submission deadline on ETS

2.  Register your organization on ETS well in advance of the important deadlines for your first tender on ETS viz ‘Date and Time of Closure of Procurement of Tender Documents’ and ‘Last Date and Time of Receipt of Bids’. Please note that even after acceptance of your registration by the Service Provider, to respond to a tender you will also require time to complete activities related to your organization, such as creation of users, assigning roles to them, etc.

3. Get your organization's concerned executives trained on ETS well in advance of your first tender submission deadline on ETS

4. Submit your bids well in advance of tender submission deadline on ETS (There could be last minute problems due to internet timeout, breakdown, et al)

5. It is the responsibility of each bidder to remember and securely store the Pass-Phrase for each Bid-Part submitted by that bidder. In the event of a bidder forgetting the Pass-Phrase before the expiry of deadline for Bid-Submission, facility is provided to the bidder to ‘Annul Previous Submission’ from the Bid-Submission Overview page and start afresh with new Pass-Phrase(s)

6. ETS will make your bid available for opening during the Online Public Tender Opening Event (TOE) ‘ONLY IF’ your ‘Status pertaining Overall Bid-Submission’ is ‘Complete’.

For your record, you can generate and save a copy of ‘Final Submission Receipt’. This receipt can be generated from 'Bid-Submission Overview Page' only if the ‘Status pertaining overall Bid-Submission’ is ‘Complete’.

NOTE :

While the first three instructions mentioned above are especially relevant to first-time users of ETS, the fourth, fifth and sixth instructions are relevant at all times.

Minimum Requirements at Bidder’s End

Computer System with good configuration (Min PIV, 1 GB RAM, Windows XP with  Service Pack 3)  Broadband connectivity  Microsoft Internet Explorer 6.0 or above  Digital Certificate(s)

Project Management Unit

National Hydrology Project (NHP)

Central Water Commission (CWC)

West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605

E-mail:- rdcdte-cwc@nic.in, Fax No+91-11-26181267,

Telephone + 91-11-26100285/26108075

# Section I.  Instructions to Proposers

Instructions to Proposers (ITP)

A.	Introduction

| 1. Scope of Proposals | 1.1 Project Management Unit, National Hydrology Project (NHP), CWC (hereinafter called the Purchaser), invites Proposals for setting up Empanelment for the supply of Goods / equipment’s (as described in the Schedule of Requirements). |
| --- | --- |
|  | 1.2 Throughout these documents, the terms “writing” means any handwritten, typewritten, or printed communication, including telex, cable, and facsimile transmission, and “day” means calendar day. Singular also means plural. |
| 2. Source of Funds | 2.1 The Government of India (hereinafter called the Borrower) has applied for a Loan from the International Bank for Reconstruction and Development (called “the Bank” in these Documents) equivalent to US$ 175 Million U.S. dollars toward the cost of the National Hydrology Project (NHP). The Borrower intends to apply a part of the proceeds of this loan to eligible payments under the Purchase Orders under Empanelment for which these documents are issued. |
|  | 2.2 Payment by the Bank will be made only at the request of the Borrower and upon approval by the Bank in accordance with the terms and conditions of the Loan Agreement, and will be subject in all respects to the terms and conditions of that Agreement. The Loan Agreement prohibits a withdrawal from the loan account for the purpose of any payment to persons or entities, or for any import of Goods, if such payment or import, to the knowledge of the Bank, is prohibited by a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations. No party other than the Borrower shall derive any rights from the Loan Agreement or have any claim to the loan proceeds. |
| 3. Fraud and Corruption | 3.1 The Bank requires that Borrowers (including beneficiaries of Bank loans), proposer, bidders, suppliers, contractors and their agents (whether declared or not), sub-contractors, sub-consultants, service providers or suppliers, and any personnel thereof, observe the highest standard of ethics during the procurement and execution of Bank-financed contracts. In pursuance of this policy, the Bank: (a) defines, for the purposes of this provision, the terms set forth below as follows: (i) “corrupt practice” is the offering, giving, receiving or soliciting, directly or indirectly, of anything of value to influence improperly the actions of another party ; (ii) “fraudulent practice” is any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation ; (iii) “collusive practice” is an arrangement between two or more parties designed to achieve an improper purpose, including to influence improperly the actions of another party ; (iv) “coercive practice” is impairing or harming, or threatening to impair or harm, directly or indirectly, any party or the property of the party to influence improperly the actions of a party ; (v) “obstructive practice” is (aa) deliberately destroying, falsifying, altering or concealing of evidence material to the investigation or making false statements to investigators in order to materially impede a Bank investigation into allegations of a corrupt, fraudulent, coercive or collusive practice; and/or threatening, harassing or intimidating any party to prevent it from disclosing its knowledge of matters relevant to the investigation or from pursuing the investigation, or (bb) acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under Clause 6 of the Empanelment. (b) will reject a Proposal for award if it determines that the proposer or bidder recommended for award, or any of its personnel, or its agents, or its sub-consultants, sub-contractors, service providers, suppliers and/or their employees, has, directly or indirectly, engaged in corrupt, fraudulent, collusive, coercive or obstructive practices in competing for the contract in question; (c) will declare mis-procurement and cancel the portion of the loan allocated to a contract if it determines at any time that representatives of the Borrower or of a recipient of any part of the proceeds of the loan engaged in corrupt, fraudulent, collusive, coercive or obstructive practices during the procurement or the implementation of the contract in question, without the Borrower having taken timely and appropriate action satisfactory to the Bank to address such practices when they occur, including by failing to inform the Bank in a timely manner at the time they knew of the practices; (d) will sanction a firm or individual, at any time, in accordance with prevailing Bank’s sanctions procedures , including by publicly declaring such firm or individual ineligible, either indefinitely or for a stated period of time: (i) to be awarded a Bank-financed contract; and (ii) to be a nominated sub-contractor, consultant, supplier or services provider of an otherwise eligible firm being awarded a Bank-financed contract; |
|  | 3.2 Furthermore, Proposers shall be aware of the provision stated in Clause 7 of the Empanelment. 3.3 In pursuance of the policy defined in ITP Sub-Clause 3.1, the Bank will cancel the portion of the loan allocated to a Contract for Goods or works if it at any time determines that corrupt or fraudulent practices were engaged in by the representatives of the Borrower or of a beneficiary of the loan during the procurement or the execution of that Contract, without the Borrower having taken timely and appropriate action satisfactory to the Bank to remedy the situation. |
| 4. Eligibility | 4.1 Except as provided in ITP Sub-Clauses 4.2 and 4.3, this Competition process is open to all firms from eligible source countries, as defined in World Bank’s “Guidelines: Procurement of goods, works and non-consulting services under IBRD loans and IDA credits & grants by World Bank borrowers” dated January 2011 revised July 2014 (called “Procurement Guidelines” in this bid document). The Bank maintains a list of countries from which Bidders, Proposers, Goods, and Services are not eligible to participate in procurement financed by the Bank. The list is regularly updated and can be obtained from the Public Information Center (PIC) of the World Bank. Joint ventures including members from ineligible source countries or including ineligible firms shall not be permitted to bid. |
|  | 4.2 Firms of a member country may be excluded from bidding if: (a) either: (i) as a matter of law or official regulation, the Borrower’s country prohibits commercial relations with that country, provided that the Bank is satisfied that such exclusion does not preclude effective competition for the supply of the Goods required; or (ii) by an act of compliance with a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations, the Borrower’s country prohibits any import of Goods from that country or any payments to persons or entities in that country. |
|  | (b) a firm (or any affiliate that directly or indirectly controls, is controlled by, or is under common control with that firm) has been engaged by (i) the Borrower or (ii) the Purchaser or (iii) End Implementing Agency that has been duly authorized to act on behalf of the Borrower or Purchaser to provide consulting services for the preparation of the design, specifications, and other documents to be used for the procurement of the Goods described in these Bidding Documents. |
|  | (c) government-owned enterprises in the Borrower’s country may participate only if they can establish that they (i) are legally and financially autonomous and (ii) operate under commercial law. No dependent agency of the Borrower or Sub-Borrower or Purchaser under a Bank-financed project shall be permitted to bid or submit a Proposal for the procurement of Goods under the project. |
|  | 4.3 A firm declared ineligible by the Bank in accordance with ITP Sub-Clause 3.1 (c) shall be ineligible to bid for a Bank-financed contract during the period of time determined by the Bank. |
|  | 4.4 Pursuant to ITP Sub-Clause 14.1, the Proposer shall furnish, as part of its Proposer, documents establishing, to the Purchaser’s satisfaction, the Proposer’s eligibility to participate in Bank financed procurement process. |
|  | 4.5 Proposers shall provide such evidence of their continued eligibility satisfactory to the Purchaser as the Purchaser shall reasonably request. |
| 5. Eligible Goods and Services | 5.1 Funds from Bank loans are disbursed only on account of expenditures for the Goods and Services, provided by nationals of, and produced in or supplied from, eligible source countries as defined in the Procurement Guidelines. The Bank maintains a list of countries from which Proposers, Bidders, Goods, and Services are not eligible to participate in procurement financed by the Bank. The list is regularly updated and can be obtained from the Public Information Center of the World Bank. Goods are produced in a Bank member country when they are mined, grown, or manufactured in the territory of that country. Goods produced or Services supplied from a Bank member country may be excluded if that member country is subject to the conditions specified in ITP Sub-Clause 4.2 (a) (i) or (ii). 5.2 For purposes of this clause, the nationality of the bidder or Proposer is distinct from the country from where the Goods and Services are supplied. 5.3 For purposes of this clause, (a) the term “Goods” includes any Goods that are the subject of this Invitation for Proposals and (b) the term “Services” includes related services such as transportation, insurance, commissioning, and training. |

B.	First Stage Competition

| 6. Documents Establishing Eligibility of Goods and Services and Conformity to Documents for setting up Empanelment | 6.1 Pursuant to ITP Clause 14, the Proposer shall furnish, as part of its Proposal, documents establishing, to the Purchaser’s satisfaction, the eligibility of the Goods and services to be supplied under the Empanelment. 6.2 The documentary evidence of the eligibility of the Goods and Services shall consist of a statement in the Proposal of the country of origin of the Goods and Services offered. |
| --- | --- |
|  | 6.3 The documentary evidence of conformity of the goods and services to the document for setting up Empanelment may be in the form of literature, drawings, and data and shall consist of: |
|  | (a) a detailed description of the essential technical and performance characteristics of the Goods; |
|  | (b) an item-by-item commentary on the Purchaser’s Technical Specifications demonstrating substantial responsiveness of the Goods and Services to those specifications, or a statement of deviations and exceptions to the provisions of the Technical Specifications; |
|  | 6.4 Wherever applicable, the Goods to be supplied under the Empanelment shall be registered (if applicable) with the relevant authority in India. A Proposer who has already registered its Goods by the time of submission of Proposal should submit a copy of the Registration Certificate with its Proposal. Otherwise, the successful Proposer, by the time of signing of Empanelment, shall submit to the Purchaser a copy of the Registration Certificate of the Goods for use in India. |
|  | 6.5 For purposes of the commentary to be furnished pursuant to ITP Clause 6.3 (b) above, the Proposer shall note that standards as well as references to brand names designated by the Purchaser in its Technical Specifications are intended to be descriptive only and not restrictive. The Proposer may substitute alternative standards, brand names, and/or catalog numbers in its Proposal, provided that it demonstrates to the Purchaser’s satisfaction that the substitutions ensure substantial equivalence to those designated in the Technical Specifications. |
| 7. Documents Establishing Qualifications of the Proposer | 7.1 The Proposer shall provide documentary evidence to establish to the Purchaser’s satisfaction that: |
|  | (a) the Proposer has the financial, technical, and production capability necessary to perform the Empanelment, meets the qualification criteria specified in the Section II, and has a successful performance history in accordance with criteria specified in the Section II. |
|  | (b) in the case of a Proposer offering to supply Goods, that the Proposer did not manufacture or otherwise produce, the Proposer has been duly authorized by the manufacturer or producer of such Goods to supply the Goods in India (as per Manufacturer’s Authorization Form in Section VI); |
|  | (c) in the case of a Proposer who is not doing business within India (or for other reasons will not itself carry out service/maintenance obligations), the Proposer is or will be (if awarded the Contract) represented by a local service/maintenance provider in India equipped and able to carry out the Proposer’s warranty obligations prescribed in the Conditions of Contract and/or Technical Specifications; and |
|  | (d) the Proposer meets the qualification criteria listed in the Section II. |
| 8. One Proposal per Proposer | 8.1 A firm shall submit only one Proposal as Individual Entity, Joint Ventures are not allowed. A firm that submits more than one Proposal will cause all the Proposals with the firm’s participation to be disqualified. |
| 9. Cost of preparation and submission of Proposals | 9.1 The Proposer shall bear all costs associated with the preparation and submission of its Proposal, and the Purchaser will in no case be responsible or liable for those costs, regardless of the conduct or outcome of the competition process. |

| 10. Content of Document for setting up Empanelment | 10.1 The Documents for setting up Empanelment are those stated below and should be read in conjunction with any addendum issued in accordance with ITP Clause 12. |
| --- | --- |
|  | Section I. Instructions to Proposers (ITP) Section II. Qualification Requirements Section III Eligible Countries Section IV. Form of Empanelment Section V Technical Specifications Section VI Sample Forms |
|  | 10.2 The “Invitation for Proposals” does not form part of the Document for setting up Empanelment and is included as a reference only. In case of discrepancies between the Invitation for Proposals and the Document for setting up Empanelment listed in 10.1 above, said Document for setting up Empanelment will take precedence. |
| 11. Clarification of Document for setting up Empanelment | 11.1 A prospective Proposer requiring any clarification of the Document for setting up Empanelment shall contact the Purchaser in writing at the Purchaser’s address : Director, River Data Directorate, Central Water Commission, West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605 (E-mail:- rdcdte-cwc@nic.in, Fax No+91-11-26181267, Telephone + 91-11-26100285 / 26108075) . The Purchaser will respond in writing to any request for clarification received no later than fourteen (14) calendar days prior to the deadline of submission of Proposals. Copies of the Purchaser’s response shall be uploaded on www.mowr.gov.in/ www.cwc.gov.in/ www.eprocure.gov.in/ www.indiawrm.org / www.tcil-india-electronictender.com. 11.2 A pre-proposal meeting will be held on 17.10.2016 at 11:00 Hrs. at the address given in sub-clause 11.1 above to provide additional clarifications. Non-attendance at the pre-proposal meeting will not be a cause for disqualification of a Proposer. Interested Proposers should depute their staff only to attend the pre-bid meeting. |
| 12. Amendment of Document for setting up Empanelment | 12.1 At any time prior to the deadline for submission of Proposals, the Purchaser may amend the Document for setting up Empanelment by issuing Addenda. |
|  | 12.2 Any addendum thus issued shall be part of the Document for setting up Empanelment pursuant to ITP Sub-Clause 10.1 and shall be communicated in writing to all purchasers of the Document for setting up Empanelment and will be binding on them. Proposers are required to immediately acknowledge receipt of any such amendment, and it will be assumed that the information contained in the amendment will have been taken into account by the Proposer in its Proposal. The addendum will also be uploaded on the website www.mowr.gov.in/ www.cwc.gov.in/ www.eprocure.gov.in/ www.indiawrm.org / www.tcil-india-electronictender.com. Proposers would be solely responsible for ensuring that any subsequent addenda issued thereafter and available in website is also downloaded / incorporated while preparing and submitting Proposals. |
|  | 12.3 To give prospective Proposers reasonable time in which to take the amendment into account in preparing their Proposals, the Purchaser shall extend, at its discretion, the deadline for submission of Proposals, in which case, the Purchaser will notify all Proposers by email confirmed in writing of the extended deadline. |
| 13. Language of Proposal | 13.1 The Proposal, as well as all correspondence and documents relating to the Proposal exchanged by the Proposer and the Purchaser, shall be written in English language. Supporting documents and printed literature furnished by the Proposer may be in another language provided they are accompanied by an accurate translation of the relevant passages in English, in which case, for purposes of interpretation of the Proposal, the translation shall govern. |
| 14. Documents Constituting the Proposal | 14.1 The Proposal submitted by the Proposer shall comprise the following: |
|  | (a) duly filled-in Form of Proposal, in accordance with the forms indicated in Section VI; (b) original form of proposal security in accordance with the provisions of ITP Sub-Clause 17 (Proposal Security); (c) written power of attorney authorizing the signatory of the Proposal to commit the Proposer; (d) the documentary evidence in accordance with ITP Sub-Clause 4.4 establishing to the Purchaser’s satisfaction the Proposer’s eligibility to submit Proposal including but not limited to documentary evidence that the Proposer is legally incorporated in a territory of an eligible source country as defined under ITP Clause 4; (e) documentary evidence establishing to the Purchaser’s satisfaction, and in accordance with ITP Clause 6 that the Goods and ancillary services to be supplied by the Proposer are eligible Goods and Services, pursuant to ITP Clause 5, and that they conform to the Document for setting up Empanelment; (f) documentary evidence establishing to the Purchaser’s satisfaction, and in accordance with ITP Clause 7 that the Proposer is qualified to perform the Contract if its Proposal is accepted.; and (h) The manufacturers’ authorization form in Form 3 of Section VI. |
| 15. Proposal Form | 15.1 The Proposer shall complete the Proposal Form and other forms furnished in the Document for setting up Empanelment, indicating the Goods to be supplied, a brief description of the Goods, their country of origin, etc. |
| 16. Period of Validity of Proposals | 16.1 Proposals shall remain valid for the period of 120 days after the date of proposal submission specified in ITP Clause 21. |
|  | 16.2 In exceptional circumstances, prior to expiry of the original proposal validity period, the Purchaser may request that the Proposers extend the period of validity for a specified additional period. The request and the responses thereto shall be made in writing. A Proposer may refuse the request without forfeiting its proposal security. A Proposer agreeing to the request will not be required or permitted to modify its Proposal, but will be required to extend the validity of its proposal security for the period of the extension. |
| 17. Proposal Participation fee | 17.1 The Proposer shall furnish, as part of its Proposal, a non- refundable proposal participation fee of INR 15000/- |
|  | 17.2 Deleted. |
|  | 17.3 The proposal security shall be denominated in Indian Rupees, and shall be in the form of demand draft in favour of the purchaser. |
|  | 17.4 Any Proposal not accompanied by an acceptable proposal participation fee shall be rejected by the Purchaser as non-responsive and not taken up for evaluation. The proposal participation fee should reach the purchaser’s office before the opening date and time of Empanelment document. |
| 17a.Empanelment fee | 17.5 The Proposer shall furnish, as part of its Proposal, a refundable empanelment fee of INR 5000/- per item. |
|  | 17.6 Empanelment fees shall remain valid for a period of 45 days beyond the validity period for the Proposal. |
|  | 17.7 The Empanelment fees shall be denominated in Indian Rupees, and shall be, at the Proposer’s option, in one of the following forms: |
|  | (a) a demand draft drawn in favour of the Purchaser; (b) a (bank) guarantee issued by a nationalized/scheduled bank in India selected by the Proposer or any reputable bank located in any eligible country. The format of the (bank) guarantee shall be in accordance with the form of proposal security included in Section VI or any other form acceptable to the Purchaser. |
|  | 17.8 Any Proposal not accompanied by an acceptable empanelment fee shall be rejected by the Purchaser as non-responsive and taken up for evaluation. The empanelment fee should reach the purchaser’s office before the opening date and time of Empanelment document. 17.9 The empanelment fee shall be refunded six month after the expiry of empanelment term. |
| 18. Alternative Proposals by Proposers | 18.1 Alternative Proposals shall not be accepted. |
| 19. Format and Signing of Proposal | 19.1 After uploading the proposal on e-procurement portal of TCIL, the Proposer shall prepare one original and 2 copies/sets of the Proposal, clearly marking each one as “ORIGINAL PROPOSAL” and “COPY OF PROPOSAL,” as appropriate. In the event of any discrepancy between them, the original shall govern. |
|  | 19.2 The original and all copies of the Proposal, each consisting of the documents listed in ITP Sub-Clause 14.1, shall be typed or written in indelible ink and shall be signed by the Proposer or a person or persons duly authorized to bind the Proposer to the Contract. The later authorization shall be indicated by written power of attorney, which pursuant to ITP Sub-Clause 14.1 (d) shall accompany the Proposal. |
| 20. Sealing and Marking of Proposals | 20.1 The Proposer shall enclose the original and 2 copies of the Proposal, in separate sealed envelopes, duly marking the envelopes as “ORIGINAL” and “COPY.” The envelopes containing the original and copies shall then be enclosed in another envelope. |
|  | 20.2 The inner and outer envelopes shall: bear the name and address of the Proposer; (b) be addressed to the Purchaser at the address below: Project Monitoring Unit, National Hydrology Project, Central Water Commission (CWC), West Block – I, 2nd Floor, Wing No – 4, Rama Krishna Puram, New Delhi 110605 E-mail:- rdcdte-cwc@nic.in Fax No+91-11-26181267, Telephone + 91-11-26100285/26108075. (c) bear the specific identification of this Competition process e.g. Proposal for setting up Empanelment, IFP number HO: CPMU:NHP/2016 -RDD/01 (d) bear a statement “DO NOT OPEN BEFORE 11.11.2016 AT 16:00 hrs”. |
|  | 20.3 If the outer envelope is not sealed and marked as required by ITP Sub-Clause 20.2, the Purchaser will assume no responsibility for the misplacement or premature opening of the Proposal. 20.4 These proposal copies shall only be used for internal official purpose only. However, the proposal shall be evaluated based on the document submitted through e-procurement portal only. |
| 21. Deadline for Submission of Proposals | 21.1 Proposals must be submitted online only at TCIL website URL https://www.tcil-india-electronictender.com before 11.11.2016 upto 15:00 hrs. Tenderer/Contractor is advised to follow the instructions provided in the 'Instructions to the Contractor/Tenderer for the e-submission of the bids online through TCIL portal. In the event of the specified date for the submission of Proposals being declared a holiday for the Purchaser, the Proposals will be received up to the appointed time on the next working day. |
|  | 21.2 The Purchaser may, at its discretion, extend the deadline for the submission of Proposals by amending the document for setting up Empanelment in accordance with ITP Sub-Clause 12.3, in which case all rights and obligations of the Purchaser and Proposers previously subject to the deadline will thereafter be subject to the deadline as extended. 21.3 Any interlineation, erasures, or overwriting to correct errors made by the Proposer should be initialed by the person or persons signing the Proposal. 21.4 The Proposer shall furnish in the Proposal Form (a sample of which is provided in the Sample Forms Section of this document) information regarding commissions or gratuities, if any, paid or to be paid to agents relating to this Proposal and to the execution of the Agreement/Purchase Order if the Proposer is awarded the Agreement/Purchase Order. |
| 22. Late Proposals | 22.1 Any Proposal received by the Purchaser after the deadline for submission of Proposals prescribed by the Purchaser in the ITP Clause 21 will be not be accepted by the e-procurement system. |
| 23. Modification and Withdrawal of Proposals | 23.1 The Proposer may modify or withdraw its Proposal after submission, provided that written notice of the modification, or withdrawal of the Proposals duly signed by an authorized representative, is received by the Purchaser prior to the deadline prescribed for submission of Proposals. No Proposal can be modified subsequent to the deadline for submission of Proposals. |
|  | 23.2 The Proposer’s modification shall be prepared, sealed, marked, and dispatched as follows: The Proposer shall provide an original and the number of copies specified ITP Sub-Clause 20.1 of any modifications to its Proposal, clearly identified as such, in two inner envelopes duly marked “PROPOSAL MODIFICATION-ORIGINAL” and “PROPOSAL MODIFICATION-COPIES.” The inner envelopes shall be sealed in an outer envelope, which shall be duly marked “PROPOSAL MODIFICATION.” (b) Other provisions concerning the marking and dispatch of Proposal modifications shall be in accordance with ITP Sub-Clauses 20.2 and 20.3. |
|  | 23.3 A Proposer wishing to withdraw its Proposal shall notify the Purchaser in writing prior to the deadline prescribed for Proposal submission. A withdrawal notice shall be received prior to the deadline for submission of Proposals. The notice of withdrawal shall: (a) be addressed to the Purchaser at the address named in the ITP Sub-Clause 20.2(b), (b) bear the specific identification of the IFP title and IFP number, and the words “PROPOSAL WITHDRAWAL NOTICE,” and (c) be accompanied by a written power of attorney authorizing the signatory of the withdrawal notice to withdraw the Proposal. |
|  | 23.4 Proposals requested to be withdrawn in accordance with ITP Sub-Clause 23.3, shall be returned unopened to the Proposers. |
|  | 23.5 No Proposal may be withdrawn after the Proposal submission deadline. Withdrawal of a Proposal after proposal submission deadline may result in the forfeiture of the Proposer’s proposal security, pursuant to ITP Sub-Clause 17.7. |
| 24. Proposal Opening | 24.1 The Purchaser will e-open all Proposals, including withdrawal notices and modifications, in public, in the presence of Proposers’ representatives who choose to attend, at 16:00 hrs on 11.11.2016 and at the place specified in the ITP Sub-Clause 20.2(b). Proposers’ representatives shall sign a register as proof of their attendance. In the event of the specified date of the proposal opening being declared a holiday for the Purchaser, the Proposals shall be opened at the appointed time and location as per corrigendum issued by the purchaser. |
|  | 24.2 Envelopes marked “WITHDRAWAL” shall be read out and the envelope with the corresponding Proposal shall not be opened but returned to the Proposer. No Proposal withdrawal notice shall be permitted unless the corresponding withdrawal notice is read out at proposal opening. Envelopes marked “MODIFICATION” shall be read out and opened with the corresponding Proposal. |
|  | 24.3 Proposals shall be opened one at a time, reading out: the name of the Proposer and whether there is a modification; the presence or absence of a proposal security; the presence or absence of requisite powers of attorney; and any other such details as the Purchaser may consider appropriate. No Proposal shall be rejected at proposal opening except for late proposals pursuant to Sub-Clause 22.1. |
|  | 24.4 Proposals (and modifications sent pursuant to ITP Sub-Clause 23.2) that are not opened and read out at proposal opening shall not be considered further for evaluation, irrespective of the circumstances. |
|  | 24.5 The Purchaser will prepare minutes of the proposal opening at the end of the opening session, including, as a minimum: the name of the Proposer and whether there was a withdrawal or modification; the presence or absence of a proposal security; the presence or absence of requisite powers of attorney. The Proposer’s representatives who are present shall be requested to sign the minutes. The omission of a Proposer’s signature on the minutes shall not invalidate the content and effect of the minutes. The minutes should be distributed to all Proposers who request them. |
| 25. Clarification of Proposals | 25.1 During evaluation of the Proposals, the Purchaser may, at its discretion, ask the Proposer for a clarification of its Proposal. The request for clarification and the response shall be in writing, and no change in the substance of the Proposal shall be sought, offered, or permitted. The maximum time limit for response from proposer on the clarification sought shall be two weeks from the date of written communication. |
| 26. Confidentiality | 26.1 Information relating to the examination, clarification, evaluation, and comparison of Proposals, and recommendations for the award of a Contract shall not be disclosed to Proposers or any other persons not officially concerned with such process until the notification of Contract award is made to all Proposers. |
|  | 26.2 Any effort by the Proposer to influence the Purchaser in the Purchaser’s proposal evaluation, proposal comparison, or contract award decisions may result in the rejection of the Proposer’s Proposal. |
|  | 26.3 From the time of proposal opening to the time of Contract award, if any Proposer wishes to contact the Purchaser on any matter related to its Proposal, it should do so in writing. |
| 27. Examination of Proposals and Determination of Responsiveness | 27.1 The Purchaser will examine the Proposals to determine whether they are complete, whether any computational errors have been made, whether required sureties have been furnished, whether the documents have been properly signed, and whether the Proposals are generally in order. |
|  | 27.2 The Purchaser may waive any minor informality, nonconformity, or irregularity in a Proposal that does not constitute a material deviation, provided such waiver does not prejudice or affect the relative ranking of any Proposer. |
|  | 27.3 Prior to the detailed evaluation, pursuant to ITP Clause 28, the Purchaser will determine whether each Proposal is of acceptable quality, is complete, and is substantially responsive to the document for setting up Empanelment. For purposes of this determination, a substantially responsive Proposal is one that conforms to all the terms, conditions, and specifications of the document for setting up Empanelment without material deviations, exceptions, objections, conditionalities, or reservations. A material deviation, exception, objection, conditionality, or reservation is one: (i) that limits in any substantial way the scope, quality, or performance of the Goods and related Services; (ii) that limits, in any substantial way that is inconsistent with the document for setting up Empanelment, the Purchaser’s rights or the successful Proposer’s obligations under the Contract; and (iii) that the acceptance of which would unfairly affect the competitive position of other Proposers who have submitted substantially responsive Proposals. |
|  | 27.4 If a Proposal is not substantially responsive, it will be rejected by the Purchaser and may not subsequently be made responsive by the Proposer by correction of the nonconformity. However, the proposer is free to modify its design to become responsive and bid for acceptance during subsequent evaluation as per clause 31 of ITP The Purchaser’s determination of a Proposal’s responsiveness is to be based on the contents of the Proposal itself without recourse to extrinsic evidence. |
| 28. Evaluation of Proposals | 28.1 The Purchaser will determine to its satisfaction whether the Proposers that are selected as having submitted the responsive Proposals are qualified to perform the Contract satisfactorily, in accordance with the criteria listed in ITP Sub-Clause 7.1 read with Section II. |
|  | 28.2 The determination will evaluate the Proposer’s financial, technical, and production capabilities. It will be based on an examination of the documentary evidence of the Proposer’s qualifications submitted by the Proposer, pursuant to ITP Sub-Clause 7.1, as well as other information the Purchaser deems necessary and appropriate. The Purchaser will also determine whether the products offered by the Proposer meet the technical specifications given in Section V. 28.3 For qualification of particular item(s)/product(s), the proposer shall submit the testing and certification reports relevant to each equipment like IMD, ISRO, WMO, UL / CE Certification. The purchaser may, at its discretion, ask the proposer to submit a quality certificate from an independent laboratory for the item/product for its testing as per standard and after satisfactory testing at Testing Lab conforming to specification and requirement, the item(s)/product(s) shall qualify for Empanelment. |
|  | 28.4 An affirmative post qualification determination will be a prerequisite for award of the Empanelment to the Proposers. A negative determination will result in rejection of the Proposer’s Proposal. 28.5 All the cost related to transport/shipment, packing of equipment for testing has to be borne by the proposer. |
| 29. Award Criteria | 29.1 The Purchaser will award the Empanelments to the Proposers whose Proposals have been determined to be substantially responsive, provided further that the Proposers are determined to be qualified to perform the Contract satisfactorily, pursuant to ITP Clause 28. |
| 30. Purchaser’s Right to Accept Any Proposal and to Reject Any or All Proposals | 30.1 The Purchaser reserves the right to accept or reject any Proposal, or to annul the first stage competition and reject all Proposals at any time prior to Empanelment award, without thereby incurring any liability to the affected Proposer or Proposers. |
| 31. Participation by new Proposers | 31.1 This being an open Empanelment, new Proposers may submit their Proposals (in the format given in original document for setting up the Empanelment) anytime after initial signing of these Agreements. Such offers received after initial signing of Empanelments will be opened on at least once on yearly basis (the date of opening will be intimated on the website of MoWR/NHP/CWC and will be evaluated on the basis of qualification and evaluation criteria specified in the original document for setting up the Empanelment. |
| 32. Notification of Award | 32.1 Prior to the expiration of the period of proposal validity, the Purchaser will notify the successful Proposers in writing by registered letter or by e-mail, to be subsequently confirmed in writing by registered letters, that their Proposals have been accepted. |
|  | 32.2 The proposer will be notified of the Empanelment. |
|  | 32.3 Upon the successful Proposers furnishing of the signed Empanelment and empanelment fee pursuant to ITP Clause 34, the Purchaser will promptly notify each unsuccessful Proposer. |
|  | 32.4 If, after notification of empanelment, a Proposer wishes to ascertain the grounds on which its Proposal was not selected, it should address its request to the Purchaser. The Purchaser will promptly respond in writing to the unsuccessful Proposer. |
| 33. Signing of Empanelment | 33.1 Promptly after the Purchaser notifies the successful Proposers that their Proposals have been accepted, the Purchaser will send the Proposers the Form for setting up the Empanelments provided in the document for setting up Empanelment, incorporating all agreements between the parties. |
|  | 33.2 Within twenty-one (21) days of receipt of the Form for setting up the Empanelment, the successful Proposers shall sign and date the Form and return it to the Purchaser. |
| 34. Performance Security | 34.1 deleted. |
|  | 34.2 Failure of the successful Proposer to comply with the requirement of ITP Clause 33 or ITP Sub-Clause 34.1 shall constitute sufficient grounds for the annulment of the empanelment. |

C. SECOND STAGE COMPETITION

| 35. Invitation of Second Stage Bid | 35.1 During the period of validity of the Empanelment, the End Implementation Agencies (EIAs) / integrators will invite the second stage financial bid. The second stage price bid will specify the quantity estimated to be procured under a particular Purchase Order, the requirement of associated services along with the delivery schedule and place of delivery. The second stage bid would be open to all suppliers / bidders who are otherwise eligible and qualified to participate. During second stage, bidders would be allowed to offer only the equipment model already accepted during first stage of Empanelment. In case an integrator wishes to quote for equipment / goods not on the empanelment list, the integrator should get the equipment / goods empaneled. |
| --- | --- |
| 36. Bid Prices | 36.1 During the second stage competition, the Supplier shall indicate on the Price Schedule, the unit prices of each item, and the total Bid price of the Goods it proposes to supply under the purchase order. The procedures laid out for second stage competition would be published by EIAs. |
| 37. Award of Purchase Orders | 37.1 Deleted |
|  | 37.2 The process of second stage competition will be repeated for each of the purchase orders to be issued during the validity of Empanelment. |

# Section II.  Qualification Requirements (Ref: ITP 7.1 (a))

The Proposer (as defined in Section-I, Instruction to bidder) should submit documentary evidence on its qualifications to perform the Contract if its bid is accepted as detailed below:

(i)	that, in the case of a Proposer offering to supply Goods under the Contract which the Proposer manufactures or otherwise produces,  that the Proposer:

(a)	is incorporated in the country of manufacture of the Goods;

(b) 	is in the continuous business of manufacturing / marketing / installing the hydro-meteorological / Geophysical / water quality equipment for at-least last five years,

(c)	has manufactured and marketed the specific goods  covered by this Bidding Document during  the last three years. In support of this, data on past performance should be submitted as per Form 4 in Section VI. The minimum quantity of goods manufactured and marketed during any one of the last 3 years is as under:

| Sl No. | Item | Minimum quantity |
| --- | --- | --- |
|  | Automatic rain gauges | 200 |
|  | Rain and snow gauges | 50 |
|  | Snow depth sensor | 50 |
|  | Shaft encoder | 200 |
|  | Radar | 200 |
|  | Ultrasonic sensor | 100 |
|  | Bubbler | 100 |
|  | Pressure transducer | 200 |
|  | Automatic weather stations | 100 |
|  | Groundwater level sensor with vent tube | 200 |
|  | Groundwater level sensor without vent tube | 200 |
|  | ADCP | 50 |
|  | GSM / GPRS modem | 300 |
|  | INSAT radio | 500 |
|  | VSAT trans-receiver | 200 |
|  | Data collection platform (data-logger) | 500 |
|  | Multi electrode resistivity imaging system | 100 |
|  | Time domain electromagnetic (tem) equipment | 100 |
|  | Resistivity meter (signal averaging system) | 100 |
|  | Multi-parameter digital geophysical logger (500 m) | 100 |
|  | Multi-parameter digital geophysical logger (1000 m) | 100 |
|  | Resistivity meter (indigenous or equivalent) | 50 |
|  | Snow water equivalent | 50 |
|  | Electromagnetic Flow Meter | 100 |
| 25) | Water Quality Sonde | 50 |

(ii)	When offering their bid for more than one item, the Proposer or the manufacturer whose product is offered by the Proposer must provide evidence that it meets or exceeds all the individual requirements for the items being applied for in regard to manufacturing and marketing capacity, as mentioned above

In case the Proposer or the manufacturer whose product is offered by the Proposer fails to fully meet any of these criteria, it will be qualified only for those items for which the Proposer meets the above requirement.

(iii)	that, in the case of a Proposer offering to supply Goods under the Contract that the Proposer does not manufacture or otherwise produce, the Proposer has been duly authorized by a manufacturer of the Goods that meets the above criteria, to supply the Goods in India, as per authorization Form 3 in Section VI; and

(iv)	The Proposer shall also furnish the following documents along with his Proposal:

(a)	copies of its audited financial statements for the past three financial years;

(b)	list of major supply contracts conducted within the last five years as per Form 4 in Section VI.

c)	Copies of satisfactory completion certificate of performances made during last three years.

v)	For qualification of particular item(s)/product(s), the proposer shall submit the testing and certification reports relevant to each equipment like IMD,, ISRO, WMO, UL / CE Certification. The purchaser may, at its discretion, ask the proposer to submit a quality certificate from an independent laboratory for the item/product for its testing as per standard and after satisfactory testing at Testing Lab conforming to specification and requirement, the item(s)/product(s) shall qualify for Empanelment.

vi)	All the cost related to transport/shipment, packing of equipment for testing has to be borne by the proposer and equipment has to be taken back by the proposer after 7 days of its testing at their own cost.

# Section III.  Eligibility for the Provision of Goods, Works, and Services in Bank-Financed Procurement

Public Information Center

Eligibility for the Provision of Goods, Works and Services in Bank-Financed Procurement

- In accordance with Para 1.8 of the Guidelines: Procurement of Goods, Works and Non-consulting Services under IBRD Loans and IDA Credits & Grants by World Bank Borrowers of January 2011 , the Bank permits firms and individuals from all countries to offer goods, work and services for Bank-financed projects.  As an exception, firms of a Country or goods manufactured in a Country may be excluded if:

Para 1.8 (a) (i): as a matter of law or official regulation, the Borrower’s Country prohibits commercial relations with that Country, provided that the Bank is satisfied that such exclusion does not preclude effective competition for the supply of the Goods or Works required, or

Para 1.8 (a) (ii): by an Act of Compliance with a Decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations, the Borrower’s Country prohibits any import of goods from that Country or any payments to persons or entities in that Country.

- For the information of borrowers and Proposers, at the present time firms, goods and services from the following countries are excluded from this bidding1:

- With reference to paragraph 1.8 (a) (i) of the Guidelines: None

- With reference to paragraph 1.8 (a) (ii) of the Guidelines: None

________________________________________________________________________

# Section IV. Form of Empanelment

Form of Empanelment

THIS Agreement (“Agreement”) is entered into this [insert starting date], by and between Project Management Unit, National Hydrology Project (“the Purchaser”) having its principal place of business at New Delhi, and [insert Supplier’s name] (“the Supplier”) having its principal office located at [insert Supplier’s address].

This Agreement is a binding contract but imposes no obligation on the Purchaser to purchase the estimated or any quantity from the Supplier.

WHEREAS, the Purchaser wishes to have the Supplier supply the Goods hereinafter referred to, and

WHEREAS, the Supplier is willing to supply these Goods,

NOW THEREFORE THE PARTIES hereby agree as follows:

| 1. Scope | The Purchaser has entered into Empanelments with one or more Suppliers for the same items in order to allow End implementing Agencies (EIAs) (acting on behalf of Purchaser) to award Purchase Order following the second stage competition under the Empanelment. The Supplier shall supply the Goods as and when the EIAs invite the second stage bid specifying the details of Goods to be delivered along with the delivery schedule and consignee’s address. Purchase Order awards shall be made to the Empanelment Suppliers offering the lowest price during the second stage competition and offering required quantity of items to be delivered as per required schedule to the required consignees. If that Supplier (who has been awarded a Purchase Order) for any reason is not able to supply the item(s) at that time, the respective EIA shall go to the Supplier offering the next lowest price(s) and offering required quantity of items to be delivered as per required schedule to the required consignees. |
| --- | --- |
| 2. Term | This Empanelment shall be valid for a period of 3 year during the period commencing [insert starting date] and continuing through [insert completion date], or any other period (maximum duration of the agreement shall not be more than 3 years) as may be subsequently agreed by the parties in writing. |
| 3. Country of Origin | All Goods and Services supplied under this Agreement shall have their origin in the countries and territories eligible under the rules of the World Bank (“the Bank”). |
| 4. Standards | The Goods supplied under this Agreement shall conform to the standards mentioned in the Technical Specifications and, when no applicable standard is mentioned, to the authoritative standards appropriate to the Goods’ country of origin. |
| 5. Use of Documents and Information | The Supplier shall not, without the Purchaser’s prior written consent, disclose the Agreement, or any provision thereof, or any specification, plan, drawing, pattern, sample, or information furnished by or on behalf of the Purchaser in connection therewith, to any person other than a person employed by the Supplier in the performance of the Agreement. |
| 6. Inspection and Audit by the Bank | The Supplier shall permit the Bank and/or persons appointed by the Bank to inspect the Supplier’s Offices and/or the accounts and records of the Supplier and its sub-contractors, agents, personnel, consultants, service providers or suppliers relating to the performance of the Agreement, and to have such accounts and records audited by auditors appointed by the Bank if required by the Bank. The Supplier’s attention is drawn to Clause 7, which provides, inter alia, that acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under this Sub-Clause constitute a prohibited practice subject to Agreement termination (as well as to a determination of ineligibility under the Procurement Guidelines). |
| 7. Fraud and Corruption | The Bank requires that Borrowers (including beneficiaries of Bank loans), bidders, suppliers, contractors and their agents (whether declared or not), sub-contractors, sub-consultants, service providers or suppliers, and any personnel thereof, observe the highest standard of ethics during the procurement and execution of Bank-financed contracts . In pursuance of this policy, the Bank: (a) defines, for the purposes of this provision, the terms set forth below as follows: (i) “corrupt practice” is the offering, giving, receiving or soliciting, directly or indirectly, of anything of value to influence improperly the actions of another party ; (ii) “fraudulent practice” is any act or omission, including a misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain a financial or other benefit or to avoid an obligation ; (iii) “collusive practice” is an arrangement between two or more parties designed to achieve an improper purpose, including to influence improperly the actions of another party ; (iv) “coercive practice” is impairing or harming, or threatening to impair or harm, directly or indirectly, any party or the property of the party to influence improperly the actions of a party ; (v) “obstructive practice” is (aa) deliberately destroying, falsifying, altering or concealing of evidence material to the investigation or making false statements to investigators in order to materially impede a Bank investigation into allegations of a corrupt, fraudulent, coercive or collusive practice; and/or threatening, harassing or intimidating any party to prevent it from disclosing its knowledge of matters relevant to the investigation or from pursuing the investigation, or (bb) acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under Clause 6 above. (b) will reject a Proposal for award if it determines that the bidder recommended for award, or any of its personnel, or its agents, or its sub-consultants, sub-contractors, service providers, suppliers and/or their employees, has, directly or indirectly, engaged in corrupt, fraudulent, collusive, coercive or obstructive practices in competing for the contract in question; (c) will declare misprocurement and cancel the portion of the loan allocated to a contract if it determines at any time that representatives of the Borrower or of a recipient of any part of the proceeds of the loan engaged in corrupt, fraudulent, collusive, coercive or obstructive practices during the procurement or the implementation of the contract in question, without the Borrower having taken timely and appropriate action satisfactory to the Bank to address such practices when they occur, including by failing to inform the Bank in a timely manner at the time they knew of the practices; (d) will sanction a firm or individual, at any time, in accordance with prevailing Bank’s sanctions procedures , including by publicly declaring such firm or individual ineligible, either indefinitely or for a stated period of time: (i) to be awarded a Bank-financed contract; and (ii) to be a nominated sub-contractor, consultant, supplier or services provider of an otherwise eligible firm being awarded a Bank-financed contract; |
| 8. Conflict of Interest | The Supplier declares that it (or any affiliate that directly or indirectly controls, is controlled by, or is under common control with Supplier) has not been engaged by (i) the Borrower or (ii) the Purchaser or (iii) the End Implementing Agency (that has been duly authorized to act on behalf of the Purchaser) to provide consulting services for the preparation of the design, specifications, and other documents to be used for the procurement of the Goods described in this Empanelment. |

| 9. Registration of Goods | If required under the Applicable Law, Goods supplied under the Agreement shall be registered for use in India. |
| --- | --- |
| 10. Patent Rights | The Supplier shall indemnify the Purchaser against all third party claims of infringement of patent, trademark, or industrial design rights arising from use of the Goods or any part thereof. |
| 11. Performance Security | The empanelment fees is required during first stage of Empanelment. The requirement of perfroamcne security would be specified in the bid document for second stage of the contract. |
| 12. Inspections and Tests | Pre dispatch inspection (to check compliance to technical specifications defined in empanelment) may be carried out by the Purchaser or an agency appointed and paid by Purchaser for this purpose. Goods shall be dispatched only after receipt of satisfactory inspection report and communication to this effect by the End Implementing Agency/ Inspection Agency. |
| 13. Packing | . Details would be in the second stage contract |
| 14. Delivery and Documents | . Details would be in the second stage contract |
| 15. Deleted |  |
| 16. Incidental Services | The incidental services required would be specified in the bid document during second stage of the agreement. The services could include but not limited to, installation, operation, maintenance and training etc. |
| 17. Warranty | The Supplier warrants that the Goods supplied under this Agreement are new, unused, of the most recent or current models and that they incorporate all recent improvements in design and materials unless provided otherwise in the Agreement. The Supplier further warrants that all Goods supplied under this Agreement shall have no defect arising from design, materials or workmanship or from any act or omission of the Supplier, that may develop under normal use of the supplied Goods. This warranty shall remain valid for Five (5) Years after the Goods have been delivered to and accepted at the final destination indicated in the Purchase Order. (b) The Purchaser shall promptly notify the Supplier in writing of any claims arising under this warranty. Upon receipt of such notice, the Supplier shall, within the period of 48 hours and with all reasonable speed, repair or replace the defective Goods or parts thereof, without cost to the Purchaser. If the Supplier, having been notified, fails to remedy the defect(s) within the period of 48 hours, the Purchaser may proceed to take such remedial action as may be necessary, at the Supplier's risk and expense and without prejudice to any other rights which the Purchaser may have against the Supplier under the Agreement. |
| 18. Payment | The method and conditions of payment to be made to the Bidder under this empanelment shall be specified in the bid document for second stage of the contract. |
| 19. Prices | Price quoted by the Suppliers in second stage shall be fixed and firm for the duration of the Purchase Order. |
| 20. Amendments | Any variation or modification to this Agreement or the Purchase Orders shall be made only by written amendment signed by the parties. |
| 21. Deleted |  |
| 22. Applicable Law | The Agreement shall be interpreted in accordance with the laws of Union of India. |
| 23. Notices | Any notice given by one party to the other pursuant to this Agreement shall be sent to the other party in writing. The Purchaser’s addresses for notice purposes is: Project Monitoring Unit, National Hydrology Project, New Delhi. The Supplier’s addresses for notice purposes is: __________________________________________ __________________________________________ Telephone No. ______________________________ Fax No. ___________________________________ |
| 24. Taxes and Duties | The Supplier shall be entirely responsible for all taxes, duties, octroi, road permits, license fees, etc., incurred until delivery of the contracted Goods to the Purchaser. |

FOR THE PURCHASER				FOR THE SUPPLIER

Signed by ___________________	Signed by ____________________

Title: ______________________	Title: ________________________

# Section V.  Technical Specifications

## Item	1	:Automatic Rain Gauge Stations

FUNCTIONAL REQUIREMENT:

- Rainfall shall be measured using the tipping bucket method and shall be able to record cumulative rainfall.

- A spout filter shall prevent ingress of insects and debris.

- IMD/WMO certification is required.

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 0 to 2500 meter |
| Sensor |
| Sensor Type | Tipping Bucket type with Reed Switch |
| Capacity | 250 mm/hour or better |
| Resolution | 0.5 mm or better |
| Accuracy (Intensity) | 2 % or better, ±2 mm |
| General Features |
| Output Interface | SDI12/ RS 485 / / 4-20 mA / Compatible with Data logger |
| Power Supply | 12 V DC or switch rated for 12 VDC |
| Material | Corrosion Resistance Metal (Stainless steel or Aluminum) |
| Enclosure | NEMA 4 |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessaries | Sensor Mounting support, cables and other accessories as required |

## Item	2	:Rain and Snow Gauge Stations

FUNCTIONAL REQUIREMENT: To measure the hourly rainfall and snowfall

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 2000 to 5000 meter |
| Sensor |
| Sensor Type | Storage Gauge with Anti-freeze system without heating |
| Capacity | 1000 mm minimum |
| Resolution | 0.5 mm or better |
| Accuracy (Intensity) | 2 % or better, ±2 mm |
| General Features |
| Output Interface | SDI12/ RS 485 / / 4-20 mA / Compatible with Data logger |
| Power Supply | 12 V DC or switch rated for 12 VDC |
| Material | Corrosion Resistance Metal (Stainless steel or Aluminum) |
| Enclosure | NEMA 4 |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessaries | Sensor Mounting support, cables and other accessories as required |

## Item	3: Snow Depth Sensor

FUNCTIONAL REQUIREMENT: To measure the depth of snow

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Units |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 2000 to 5000 meter |
| Sensor |
| Sensor Type | Ultrasonic sensor |
| Range | 0-10 meter |
| Resolution | 1 mm or better |
| Accuracy | 0.25 % of measuring distance |
| General Features |
| Output Interface | SDI12/ RS 485 / / 4-20 mA / Compatible with Data logger |
| Power Supply | 9-18 V DC |
| Material | Corrosion Resistance Metal (Stainless steel/ Aluminum or PVC) |
| Enclosure | NEMA 4 |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessories | Sensor Mounting support, cables and other accessories as required |

## Item	4	:Shaft Encoder

FUNCTIONAL REQUIREMENT: To measure the water level

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 0 to 2500 meter |
| Sensor |
| Sensor Type | Shaft Encoder based rotary position sensor with Digital Display |
| Range | 1-100 meter |
| Resolution | 3 mm or better |
| Accuracy | 0.025 % FSO |
| Output Interface | SDI-12 / RS 485 / 4-20 mA / compatible with data logger |
| Power Supply | 12 V DC or Switch rated for 12 V DC |
| General Features |
| Material | Corrosion Resistance Metal (Stainless steel or Aluminum) |
| Enclosure | Lockable (key) box provided by the supplier to be mounted in Stilling well or Gauge hut, with IP65 or NEMA 4 protection |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Graduated Tape | The tape should be of high quality to withstand harsh and humid environment, should not get twisted or wrinkeled while operation. |
| Accessaries | Sensor Mounting support, Floats, graduated tapes (metric), wheel, counterweight, and cabling |

## Item	5	: Radar

FUNCTIONAL REQUIREMENT: To measure the water level

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | -20°C to +60°C |
| Humidity | 0 to 100 % |
| Altitude | 0 to 2500 meter |
| Sensor |
| Sensor Type | Microwave non-contact sensor |
| Range | 15M / 20M/35M/75M |
| Resolution | 3 mm or better |
| Accuracy | 0.02 % FSO |
| Beam Angle: | ≤ 16 ° |
| Output Interface | SDI-12 / RS 485 / 4-20 mA / compatible with data logger |
| Power Supply | 10-15 V DC |
| General Features |
| Material | Corrosion Resistance Metal (Stainless steel / Aluminum or PVC) |
| Enclosure | The Sensor shall be easy to dismount and replace in the event of malfunction. |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessories | Sensor Mounting support, cables and other accessories as required |
| Protection | NEMA 4 or IP65 |
| Horizontal Mounting /Installation Arrangements | Above FRL, Below a bridge girder wherever available otherwise horizontal cantilever arrangement from a mast/wall/pedastal |
| Radar Sensor should have display feature for diagnostic purpose |

## Item	6	: Ultrasonic Sensor

FUNCTIONAL REQUIREMENT: To measure the water level

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 0 to 2500 meter |
| Sensor |
| Sensor Type | Ultrasonic non-contact sensor |
| Range | Upto 10 meter |
| Resolution | 3 mm or better |
| Accuracy | 0.02 % FSO |
| Output Interface | SDI-12 / RS 485 / 4-20 mA / compatible with data logger |
| Power Supply | 10-15 V DC |
| General Features |
| Material | Corrosion Resistance Metal (Stainless steel / Aluminum or PVC) |
| Enclosure | The Sensor shall be easy to dismount and replace in the event of malfunction. |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessories | Sensor Mounting support, cables and other accessories as required |
| Protection | NEMA 4 or IP65 |

## Item	7	: Bubbler

FUNCTIONAL REQUIREMENT: To measure the water level

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 0 to 2500 meter |
| Sensor |
| Sensor Type | Continuous bubbling system and non-submersible transducer |
| Range | 15/30 PSI |
| Resolution | 0.0001 psi or better |
| Accuracy | 0.1 % FSO |
| Output Interface | SDI-12 / 4-20 mA / RS485, compatible with Data logger |
| Power Supply | 11 to 15 V DC |
| Average current Draw | <15mA based on 1 bubble per second |
| Purge | Manual line purge |
| Bubble Rate | Programmable 30–120 bubbles per minute |
| Desiccators | The bubbling mechanism and the non-submersible transducer must be equipped with a desiccating system to keep system from malfunction for a period not less than one year. |
| General Features |
| Tools | Complete tool kit for installation and routine maintenance |
| Manuals | Full documentation and maintenance instructions in English |
| Accessories | Sensor Mounting support, cables and other accessories as required |
| (*) Enclosure | NEMA4 or IP65 |

## Item	8	: Pressure Transducer

FUNCTIONAL REQUIREMENT: To measure the water level

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 0 to 2500 meter |
| Sensor |
| Sensor Type | Pressure Sensor |
| Range | Upto 30 meter of water column |
| Resolution | 3 mm or better |
| Accuracy | 0.02 % FSO |
| Output Interface | SDI-12 / RS 485 / 4-20 mA / compatible with data logger |
| Power Supply | 10-15 V DC |
| General Features |
| Material | Corrosion Resistance Metal (Stainless steel / Aluminum or PVC) |
| Enclosure | The Sensor shall be easy to dismount and replace in the event of malfunction. |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessories | Sensor Mounting support, cables and other accessories as required |
| Protection | NEMA 4 or IP65 |

## Item	9	: Automatic Weather Stations

FUNCTIONAL REQUIREMENT: To measure the weather parameter

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 0 to 2500 meter |
| Air Temperature Sensor |
| Sensor Type | Platinum resistance or better or equivalent |
| Range | -20 Degree Celsius to + 60 Degree Celsius |
| Resolution | ± 0.1°C |
| Accuracy | Within ± 0.1o Celsius in the entire working range |
| Response time | 10 Secs or lesser |
| Self-aspirated | To ensure continuous supply of air.Free from turbulence, water droplets and Radiation |
| Power Supply | 12 V DC or switch rated for 12 VDC |
| Accessories | All accessories for mounting the instrument e.g. special cross arm clamps or flag, if any, shall be provided. |
| Relative humidity Sensor |
| Sensor Type | Capacitive/ Solid State Humidity Sensor |
| Range | 0 to 100 % |
| Resolution | 1% |
| Accuracy | ±3% or better |
| Power Supply | 12 V DC or switch rated for 12 VDC |
| Response time | 10 Secs or lesser |
| Wind Speed and Direction Sensor |
| Sensor Type | Ultrasonic sensor (No moving Parts) |
| Range | 0-60 m/s for speed & 0-360 degrees for direction or better |
| Resolution | 0.1 m/s for Speed; +5 degree for direction |
| Accuracy | Better than 1% full scale |
| Response time | Less than 1 second lag in operating range |
| Mounting | All accessories for mounting the instrument e.g. special cross arm clamps or flag if any shall be provided. |
| Air Pressure Sensor |
| Sensor Type | Temperature Compensated |
| Range | 800 - 1200 hPa |
| Resolution | ± 0.01 hPa |
| Accuracy | ± 0.5 hPa |
| Power Supply | 12 V DC or switch rated for 12 VDC |
| Solar Radiation Sensor |
| Sensor Type | ISO Class 1 Pyranometer (CMP 11 or better) |
| Threshold | 120 W/m2 of direct solar irradiance |
| Methodology | Alternate shading of sensor to account for sky radiation |
| Spectral Range | 400nm to 1100 nm |
| Range | 0-2000 W/Square meter |
| Resolution | 1 W/Square meter |
| Accuracy (Including Temperature Compensation) | 3% or better |
| General Features |
| Material | Corrosion Resistance Metal (Stainless steel/ Aluminum or PVC) |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessories | Sensor Mounting support, cables and other accessories as required |
| Output Interface | SDI 12/RS 485/ 4-20 mA/ Compatible with Data logger |
| Evaporation- Pan Specification |
| Operating temperature | -5 to 60 degree Celsius. |
| Diameter of the pan | 1.2 m or more |
| Accuracy | +/- 1% |
| Accessories | As required for complete installation of the sensors and Equipment |
| Material | Clean cast seamless acrylic plastic tubing or brass sheet |
| Platform | Rot resistant timber treated with creosote or other effective Wood preservative. |
| Graduation | in milliimetre |

Item	10	: Groundwater Level Recorder (DWLR) Without Vent Tube

FUNCTIONAL REQUIREMENT: To measure and transmit the ground water level

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5-100 % |
| Altitude | 0-2500 meter |
| Sensor |
| Sensor Type | Submersible pressure transducer having vent tube, with atmospheric pressure and temperature compensation |
| Range | 30 psi The full scale water fluctuation measuring range will be specified by the implementing agency depending on the requirement 0-5/10/20/30/50 m of water column |
| Installation Depth | The installation depth will be specified by the implementing Agency depending upon their requirements 0-10/20/50/100/150/200/300m |
| Accuracy | 0.1% FSO |
| Temperature Coefficient | <0.01% Full scale/degree centigrade for water temperatures between 10˚C and 40˚C |
| Resolution | 3 mm |
| Reproducibility | 0.1% full scale or better |
| Long Term Stability | 0.1% Full scale and should ensure long term stability without any field calibration requirements except barometric compensation |
| Temperature Measuring Range | 0 to 50˚C |
| Accuracy | Better than 0.1˚C |
| Burst Pressure | >=3 Time Full scale |
| Overload Pressure | 2 Time full scale without effect on calibration |
| Over voltage protection on supply & sensor wires | Over voltage protection should be provided on power supply lines |
| Non-Vented Cable | Includes barometric sensor for post-processing |
| Output | SDI-12, RS-485 or compatible with data logger included |
| Datalogger |
| Atmospheric Pressure correction | Should be applied automatically |
| Resolution of Measurement | 1mm or better |
| Measuring interval and measuring modes | Should be programmed to store data from 1 minute one reading to 24 hours one reading with future start option. |
| Settling up Time | <30 minutes after submersion |
| Recording Capacity | Non-Volatile flash data storage of more than 1,00,000 data points (at least) |
| Memory Type | Non-Volatile memory |
| Power Supply | Should be equipped with lithium or alkaline battery pack, giving at least 5 years operation (with one transmission and four recordings per day). Battery must be replaceable in the field or in local offices of the implementing Agency or supplier. Replacement batteries must be readily available in India. |
| Battery Voltage Monitoring | Monitoring and transmission of Battery Voltage level |
| Datalogger Location | If Data logger and transmitter are integral parts of sensor, it should be located on top (near ground surface) instead of bottom |
| Built in clock | Accurate to ±1 minute per year |
| Displayed Time Resolution | 1 second |
| Over-voltage Protection | Should include lightening, over-voltage and surge protection |
| Enclosure |  |
| Enclosure for Pressure sensor and Data Logger | Data Logger should be concealed into a single tubular enclosure which is waterproof and corrosion proof. |
| Dimension | Outer diameter of sensor unit: <=50mm, (for sensor & logger only) |
| Material | Titanium, stainless Steel or other corrosion resistant material |
| Installation | The system should be provided with a suspension bracket or Well Cap allowing secure installation within the Piezometers’ headwork, including appropriate cable mounting to allow the sensor to be adjusted to the required depth |
| Direct Read Cable | The cable shall have following features: Diameter of cable should be less than 30mm Strength members for good longitudinal stability of cable The cable and contacts should be fixed or quick connect |
| Protection | IP67 with Impact Resistant |
| Communication Interface |  |
| Computer Interface | The Logger must be capable of connection to a computer via USB 2.0/USB 3.0 and supply should include the necessary interface cables. |
| Wireless Communication | Option for Bluetooth/IR/Wi-Fi interface (atleast any one of the three options specified) should be available. |
| File Format | The format of the data downloaded by communication interface shall be in standard ASCII/CSV/XML format. |
| GSM / GPRS Transmitter |
| Transmission System | GPRS/edge based data transmission system |
| Performance | Data Reception availability of 95% or better |
| Communication Direction | Utilize GPRS network for two-way TCP/IP (INTERNET) connection |
| VPN protocol | Radio to utilize VPN protocol |
| Transmission trigger | Data collection to be triggered by interrogation from Data Center, or by event based transmission triggered by remote site |
| Power Saving | Ability to disable interrogation system in order to save power at remote site |
| Communication Protocol | Data transmission to execute HTTP Post or FTPS to transmit data to the Data Center |
| Accessories | All associated equipment, including Antenna all cables and mounting hardware |
| Software |
| Operating System | Windows software for system configuration, transfer and analysis of data to computer |
| Version | English language version |
| License | All required licenses included |
| General Features |
| Battery | The battery should be easy to replace, and easily available in the market |
| Tools | Complete tool kit for installation and routine maintenance |
| Manuals | Full documentation and maintenance instructions in English |

## Item	11	: Groundwater Level Recorder (DWLR) With Vent Tube

FUNCTIONAL REQUIREMENT: To measure and transmit the ground water level

DESIGN  REQUIREMENTS: The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +60 |
| Humidity | 5-100 % |
| Altitude | 0-2500 meter |
| Sensor |
| Sensor Type | Submersible pressure transducer without vent tube, having atmospheric pressure compensation sensor on each individual equipment |
| Range | 30 psi The full scale water fluctuation measuring range will be specified by the implementing agency depending on the requirement 0-5/10/20/30/50 m of water column |
| Installation Depth | The installation depth will be specified by the implementing Agency depending upon their requirements 0-10/20/50/100/150/200/300m as specified |
| Accuracy | 0.1% FSO |
| Temperature Coefficient | <0.01% Full scale/degree centigrade for water temperatures between 10˚C and 40˚C |
| Resolution | 3 mm |
| Reproducibility | 0.1% full scale or better |
| Long Term Stability | 0.1% Full scale and should ensure long term stability without any field calibration requirements except barometric compensation |
| Temperature Measuring Range | 0 to 50˚C |
| Accuracy | Better than 0.1˚C |
| Burst Pressure | >=3 Time Full scale |
| Overload Pressure | 2 Time full scale without effect on calibration |
| Over voltage protection on supply & sensor wires | Over voltage protection should be provided on power supply lines |
| Non-Vented Cable | Includes barometric sensor for post-processing |
| Output | SDI-12, RS-485 |
| Datalogger |
| Atmospheric Pressure correction | Should be applied automatically |
| Resolution of Measurement | 1mm or better |
| Measuring interval and measuring modes | Should be programmed to store data from 1 minute one reading to 24 hours one reading with future start option. |
| Settling up Time | <30 minutes after submersion |
| Recording Capacity | Non-Volatile flash data storage of more than 1,00,000 data points (at least) |
| Memory Type | Non-Volatile memory |
| Power Supply | Should be equipped with lithium or alkaline battery pack, giving at least 5 years operation (with one transmission and four recordings per day). Battery must be replaceable in the field or in local offices of the implementing Agency or supplier. Replacement batteries must be readily available in India. |
| Battery Voltage Monitoring | Monitoring and transmission of Battery Voltage level |
| Datalogger Location | If Data logger and transmitter are integral parts of sensor, it should be located on top (near ground surface) instead of bottom |
| Built in clock | Accurate to ±1 minute per year |
| Displayed Time Resolution | 1 second |
| Over-voltage Protection | Should include lightening, over-voltage and surge protection |
| Enclosure |  |
| Enclosure for Pressure sensor and Data Logger | Data Logger should be concealed into a single tubular enclosure which is waterproof and corrosion proof. |
| Dimension | Outer diameter of sensor unit: <=50mm, (for sensor & logger only) |
| Material | Titanium, stainless Steel or other corrosion resistant material |
| Installation | The system should be provided with a suspension bracket or Well Cap allowing secure installation within the Piezometers’ headwork, including appropriate cable mounting to allow the sensor to be adjusted to the required depth |
| Direct Read Cable | The cable shall have following features: Diameter of cable should be less than 30mm Strength members for good longitudinal stability of cable The cable and contacts should be fixed or quick connect |
| Protection | IP67 with Impact Resistant |
| Communication Interface |  |
| Computer Interface | The Logger must be capable of connection to a computer via USB 2.0/USB 3.0 and supply should include the necessary interface cables. |
| Wireless Communication | Option for Bluetooth/IR/Wi-Fi interface (atleast any one of the three options specified) should be available. |
| File Format | The format of the data downloaded by communication interface shall be in standard ASCII/CSV/XML format. |
| GSM / GPRS Transmitter |
| Transmission System | GPRS/edge based data transmission system |
| Performance | Data Reception availability of 95% or better |
| Communication Direction | Utilize GPRS network for two-way TCP/IP (INTERNET) connection |
| VPN protocol | Radio to utilize VPN protocol |
| Transmission trigger | Data collection to be triggered by interrogation from Data Center, or by event based transmission triggered by remote site |
| Power Saving | Ability to disable interrogation system in order to save power at remote site |
| Communication Protocol | Data transmission to execute HTTP Post or FTPS to transmit data to the Data Center |
| Accessories | All associated equipment, including Antenna all cables and mounting hardware |
| Software |
| Operating System | Windows software for system configuration, transfer and analysis of data to computer |
| Version | English language version |
| License | All required licenses included |
| General Features |
| Battery | The battery should be easy to replace, and easily available in the market |
| Tools | Complete tool kit for installation and routine maintenance |
| Manuals | Full documentation and maintenance instructions in English |

## Item	12: ADCP

FUNCTIONAL REQUIREMENT: To measure the discharge

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | -5 to 45 Degree C |
| Humidity | 5-100 % |
| Altitude | 0-2500 meter |
| mode of operation | real time from a sailing boat/Bridge/cableway |
| Sensor |
| ADCP Type | Down looking ADCP for measurement of discharge in open channel environment |
| Velocity Profiling Range | 0.1 to 5 meter / 0.4–25 meter / 0.4 to 40 meter (Actual requirement would be specified by implementing agency based on site conditions) |
| Profiling Velocity | +/-20 m/s |
| Velocity Accuracy | 0.25% of measured velocity |
| Velocity Resolution | 0.001m/s |
| Depth Range | 0.3-80 m |
| Depth Accuracy | 1%. |
| Depth Resolution | 0.001 m |
| Positioning | Optional capability to acquire position by bottom tracking or DGPS. |
| Computations | All performed internally or on Windows-based software (also to be supplied) |
| Accessories |
| Platform | Floating platform/ Trimaran for ADCP deployment (optional as per requirement of Intendter) |
| Positioning | DGPS for positioning in case of moving bed |
| Tethers | All necessary tethers and taglines |
| Software | Windows-based software for display of velocity, discharge, depth, and width information in real-time. |
| General Features |
| Tools | Complete tool kit for installation and routine maintenance |
| Manuals | Full documentation and maintenance instructions in English |

## Item	13	: GSM / GPRS  Modem

FUNCTIONAL REQUIREMENT: To transmit data

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Ambient Site Conditions |  |
| Operating Temperature | From -20 to +60 |
| Performance | Data Reception availability of 95% or better |
| Form factor | The Transmitter should either be integral part of data logger specified above, or it should be supplied as independent unit compatible with supplied data logger |
| Specific Features |
| Communication Direction | Utilize GPRS network for two-way TCP/IP (INTERNET) connection |
| VPN protocol | Radio to utilize VPN protocol |
| Transmission trigger | Data collection to be triggered by interrogation from Data Center, or by event based transmission triggered by remote site |
| Power Saving | Ability to disable interrogation system in order to save power at remote site |
| Communication Protocol | Data transmission to execute HTTP Post or FTPS to transmit data to the Data Center |
| Accessories | All associated equipment, including Antenna all cables and mounting hardware |
| Antenna features |
| Frequency range | 900 MHz: 824-960 MHz/1800MHz:1710-1880 MHz |
| Impedance | 50 ohms |
| VSWR | ≤ 2.0 |
| Radiation | Omni-directional |
| Operating temperature | -10 to + 60 degrees Celsius |
| Connector | SMA adaptable to GSM/GPRS modem |
| Cable length | As required |

## Item	14	: INSAT Radio

FUNCTIONAL REQUIREMENT: To transmit data

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Operating Temperature | From -20 to +60 |
| Environment Relative Humidity | 0 to 100 % |
| Career Frequency | 402 - 403 MHz |
| Carrier Settability | In steps of 100 Hz from 402.0 MHz to 403.0 MHz |  |
| Modulator | PCM/BPSK |
| Data coding | NRZ(L) |
| Output Power | 3-10 W, user settable |
| Data Bit Rate | 4.8 kbps |
| Frequency Stability |  |
| Long term | Transmit frequency inaccuracy including aging of oscillator should not exceed ( 400 Hz per year. Oscillator/synthesizer should have provision to adjust for the long term drift |
| for temperature | ( 1 ppm or better (-40 to +55oC) |
| Signal Bandwidth | 6.0 KHz maximum or better |  |
| Output Power | 3-10 W (settable) |  |
| Power Stability | (1 dB |
| Spurious | -60 dB or better |
| Harmonics | -40 dB or better |  |
| Antenna cable | LMR 400 grade or better |
| Performance | Data Reception availability of 99% or better |
| Form factor | The Transmitter should either be integral part of data logger specified above, or it should be supplied as independent unit compatible with supplied data logger |
| Operating power | Switched 12V D.C controlled by data logger. |  |
| Yagi Antenna |
| Polarization | LHCP or RHCP, switchable in field |
| Gain | Minimum 11 dbi or better |
| Center Frequency | 402-403 MHz |
| Mounting | Proper mounting and Pointing arrangement for 360 degree azimuth and elevation adjustment |
| Operating Wind speed | 250 kmph |
| Wind Survival | 300 kmph |
| Material | Rust-proof and Oxidation-proof |
| Specific Features |
| Satellite System | INSAT Radio System to be Used on the INSAT Satellite operated by ISRO |
| Certification | Certificate of acceptance required by ISRO and/or IMD as part of the bid package |
| Demonstration in India | Demonstrated use of the satellite radio with at least 200 radios in current operation in India using INSAT |
| Accessories | All associated equipment, including GPS, GPS Antenna, INSAT Antenna, all cables and mounting hardware |

## Item	15	: VSAT Trans-receiver

FUNCTIONAL REQUIREMENT: To transmit and receive data

DESIGN  REQUIREMENTS:The equipment offered should conform to the following technical Specifications:

| Feature | Value |
| --- | --- |
| Operating Temperature | From -20 to +60 |
| Antenna cable | LMR 400 grade or better |
| Performance | Data Reception availability of 99% or better |
| Specific Features |
| Communication Direction | VSAT Radio system to allow two-way communication system between Data Center and remote station |
| Single Hop or double hop | Provision to use either single hop (leased lines between user and service provider hub) or double hop (via vsat) for receiving data at user end |
| Frequency Band | C Band or Extended C band (Ku or Ka band would be acceptable) |
| Bandwidth Sharing | VSAT bandwidth will be able to be shared among all stations using TDMA mode |
| Alarm Conditions | VSAT remote stations shall be able to transmit based on alarm conditions at the remote site such as critical water level or exceptional precipitation events |
| Accessories | All associated equipment, including Antenna all cables and mounting hardware |

## Item	16	: Data Collection Platform

FUNCTIONAL REQUIREMENT:

    - The DCP shall also continuously monitor the status of the instruments, power supply and communications.  In the event of failure of an instrument or disruption of any of the power sources, an alarm shall be sent back to the ERS/modeling center.

    - The sensor's signal conditioning unit should be an integral part of the system.

    - The system shall have provision to easily include and change the following information in field as mandatory requirements:

- Unique station identification code.

- Time of observation.

- Sensor identification.

    - The system shall have an integrated microprocessor based data acquisition and storage system having adequate hardware configuration and software support to serve as an interface between sensors and the communication link to perform tasks as stated below.

    - Providing necessary electrical power to the sensors and conversion of electrical output signals from the sensors into engineering values based on calibration equations stored in the memory.  Full compatibility with all types of sensors provided in the packages shall be mandatory.

    - The system should be stand-alone and all programming functions/ set-ups to be carried out through system keypad and display independent of a PC/ Laptop.

    - The system should be capable of continuous updating of the values of sensed parameters and post processing the instantaneous values into average values over a specified period of time for transmission to the DCP with earth receiving station.

    - The system shall have in-built sensor simulation system options to conduct tests on the system for field installation, two-point calibration/ re-calibration and maintenance of the sensors.

    - The system shall support the following functions:

- Easy programming set up.

- Multi tasking capability

- User friendly software programming.

- The system shall have self-diagnostic facility and be capable of displaying Station ID/ Sensor ID codes and messages on the display panel for general identification of the fault.  It should have facility to monitor these codes and other health status through an external lap top/ PC.

- Setup shall be organised in a tree of menus and sub-menus. Protection of setup parameters and data through password should be supported by the system.  In addition, the DCP shall support the manual entry of data through keypad and its display.

- Data including the setup and program files shall be transferable from the system via a serial port to PC and SD card or other suitable memory device and vice versa.

    - The data logger shall be programmable locally in field via laptop/ PC.

    - The surge suppression in form of fuse or other appropriate device shall be provided for all interfaces to protect the data logger from surges emanating from the sensors.

### 16A. Specification of Data Logger for 1-2 Sensors

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +50 Degree C |
| Humidity | 5 to 100 % |
| Altitude | 0 to 5000 meter |
| Sensor Interface |
| Analogue Inputs | 1 Analogue Input Channels |
| Analog inputs | 4 to 20 mA ; 100% over-range withstand |
| SDI Port | One SDI-12 Interface port |
| Digital Inputs | 1 Digital Channels, bidirectional |
| Pulse Input | 1 Input for Rain Gauge impulse |
| Input - Output Interfaces |
| Data Transfer | USB stick option for Data transfer |
| Port for Configuration | One Serial Port (RS232) for communication with Laptop for programming |
| 1Port for Telemetry | 2 Ports for Communication with Telemetry (GSM / VSAT / INSAT) Device (See Note 1 Below) |
| 2Display Port | Optional port for connecting external display screen for Data in running text (See Note 2 Below) |
| Computer Software |
| Operating System | Windows software for system configuration / communication |
| Version | English language version |
| Licenses | All required licenses included |
| Analog to digital converter |
| Resolution | 16 bit or better |
| Conversion Accuracy | ± 1 LSB |
| Sample Intervals | 1 sec. to 24 hr. in 1 second increments (user selectable) |
| General Features |
| Flash memory | Non-volatile Flash memory that can one store one year of data and expandable to a minimum of 1GB. |
| Resolution | A/D resolution ≥16 bit |
| Recording Interval | Individual recording intervals for each sensor/parameter |
| Firmware Operating System | Multi-tasking operating system - must log data and transmit at same time |
| Display | Inbuilt Digital Display for viewing current data and setting values |
| Power Supply | Power supply 12V DC, low current drain (quiescent ≤10.0mA) |
| Battery Voltage | Monitoring of battery voltage level |
| Internal battery | Internal battery backup for clock, Lithium Battery, storage: 2 years |
| Charge controller | Internal or External |
| User Permissions | Different user levels, system of user rights / passwords, access restricted to authorized personnel |
| Internal clock | Internal clock with drift less than 2 seconds per year or using GPS |
| Keypad | Keypad for displaying or transferring data to memory stick, configuration of data-logger and sensors |
| Real-Time Clock | GPS synchronised |
| System integrity | System integrity check procedures |
| Enclosure | for wall-mounting in a shelter / enclosure with IP65 (NEMA 4) protection or better |
| Accessories | Serial cable + adaptor (if required) for notebook connection. All accessories (fixing units, etc.) as required |
| Tools | complete tool kit for installation and routine maintenance giving full detail( number of pieces and type) |
| Manuals | full documentation and maintenance instructions in English (1 copy per station). |

Note 1: The datalogger should have atleast 2 ports for data transmission via telemetry devices (GSM / VSAT / INSAT). Both telemetry systems should work simultaneously for redundancy. The type of port required for telemetry device may be different (Serial, RS 485, RS 232, RJ-45 etc) and proposer may offer multiple models having different combination of ports.

Note 2: The port for attaching external display device to show data as running text is optional. The proposer may offer two different models, with or without port for display device.

### 16B. Specification of Data Logger for more than 2 Sensors

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -20 to +50 Degree C |
| Humidity | 5 to 100 % |
| Altitude | 0 to 5000 meter |
| Sensor Interface |
| Analogue Inputs | 8 Analogue Input Channels |
| Analog inputs | 4 to 20 mA ; 100% over-range withstand |
| SDI Port | One SDI-12 Interface port |
| Digital Inputs | 6 Digital Channels, bidirectional |
| Pulse Input | 2 Input for Rain Gauge impulse |
| Input - Output Interfaces |
| Data Transfer | USB stick option for Data transfer |
| Port for Configuration | One Serial Port (RS232) for communication with Laptop for programming |
| 1Port for Telemetry | 2 Ports for Communication with Telemetry (GSM / VSAT / INSAT) Device (See Note 1 Below) |
| 2Display Port | Optional port for connecting external display screen for Data in running text (See Note 2 Below) |
| Computer Software |
| Operating System | Windows software for system configuration / communication |
| Version | English language version |
| Licenses | All required licenses included |
| Analog to digital converter |
| Resolution | 16 bit or better |
| Conversion Accuracy | ± 1 LSB |
| Sample Intervals | 1 sec. to 24 hr. in 1 second increments (user selectable) |
| General Features |
| Flash memory | Non-volatile Flash memory that can one store one year of data and expandable to a minimum of 1GB. |
| Resolution | A/D resolution ≥16 bit |
| Recording Interval | Individual recording intervals for each sensor/parameter |
| Firmware Operating System | Multi-tasking operating system - must log data and transmit at same time |
| Display | Inbuilt Digital Display for viewing current data and setting values |
| Power Supply | Power supply 12V DC, low current drain (quiescent ≤10.0mA) |
| Battery Voltage | Monitoring of battery voltage level |
| Internal battery | Internal battery backup for clock, Lithium Battery, storage: 2 years |
| Charge controller | Internal or External |
| User Permissions | Different user levels, system of user rights / passwords, access restricted to authorized personnel |
| Internal clock | Internal clock with drift less than 2 seconds per year or using GPS |
| Keypad | Keypad for displaying or transferring data to memory stick, configuration of data-logger and sensors |
| Real-Time Clock | GPS synchronised |
| System integrity | System integrity check procedures |
| Enclosure | for wall-mounting in a shelter / enclosure with IP65 (NEMA 4) protection or better |
| Accessories | Serial cable + adaptor (if required) for notebook connection. All accessories (fixing units, etc.) as required |
| Tools | complete tool kit for installation and routine maintenance giving full detail( number of pieces and type) |
| Manuals | full documentation and maintenance instructions in English (1 copy per station). |

Note 1: The datalogger should have atleast 2 ports for data transmission via telemetry devices (GSM / VSAT / INSAT). Both telemetry systems should work simultaneously for redundancy. The type of port required for telemetry device may be different (Serial, RS 485, RS 232, RJ-45 etc) and proposer may offer multiple models having different combination of ports.

Note 2: The port for attaching external display device to show data as running text is optional. The proposer may offer two different models, with or without port for display device.

### 16C. Power Supply for DCP

The equipment offered should conform to the following technical Specifications:

| Feature | Units |
| --- | --- |
| Battery |
| Voltage | From -20 to +60 |
| Type | Sealed Maintenance free |
| Capacity | Based on site conditions and Telemetry method, to provide 21 days of backup |
| Solar Panels |
| Size | Based on Site conditions and Telemetry method used for 21 days of backup |
| Mounts | The mounts should be sturdy in design; the solar panel should not move or rotate with wind. It should have provision to adjust direction and elevation during installation for optimal solar power generation |
| Charger | Smart solar charger with protection |
| General |
| The supplier should determine optimal size of solar panels and batteries, such that system should be operational for at least 21 days in the absence of charging |

## Item 17: Multi Electrode Resistivity Imaging System

Multi- electrode Resistivity imaging with SP and IP system fully automatic, capable of recording, storing Resistivity and induced polarization data for arbitrary electrodes configuration. Software fully automatic and data acquisition, processing and interpretation is required with the facility of 2/3 – Dimensional   Resistivity imaging. The equipment should be complete with all accessories: electrodes, Multi core cable, Charger for internal batteries transportation case and other required accessories for field survey. All the equipments and accessories should be fitted in a suitable AC vehicle so that data acquisition and processing of field data can be done at the site itself. It will also provide the safety of equipments which contains various electronic circuits.

| Feature | Specification |
| --- | --- |
| General Features |
| General | The unit should be compact and light with display, main processing unit; internal memory and multi-electrode system are integrated in the same housing. No separate electronics other than cables. |
| Environmental | Weather proof, Shock proof |
| Operating Temp | From -5 Degree C to +50 Degree C |
| Interface | Equipment can be interfaced with PC or laptop for upload and load of sequence file and data file. |
| Power source | Internal and external rechargeable battery source for imaging, data acquisition and data transfer to external PC or Laptop. |
| Specific Features |
| Injection Current | Automatic injection current ranging and stacking and averaging. |
| Resistivity Measurement | Direct measurement of Resistivity, SP and IP. |
| Cable | 120 takes out in 10m spacing |
| No. of Channels | 10 channels or more. |
| Automatic Ranging | Microprocessor controlled. |
| Current measurement precession | 0.2 percent. |
| Voltage measurement precision/ resolution | 0.2 percent typical/ 1 micro volt or better |
| Noise Reduction | Continuous stacking selectable manually or automatic. |
| SP compensation | Through automatic line at drift correction. |
| Resistivity accuracy | 0.2 percent typical or better. |
| Induced Polarization | Arbitrary windows flexibility, configured to power line frequencies |
| Output |
| Output Voltage | 400 V or more. (Peak to Peak 800 V or more) |
| Output current | 2 A or more |
| Output Power | 200W or more |
| Input |
| Input Impedance | 100 Mega Ohm |
| Input Voltage protection | up to 1000V |
| External Transmitter |  |
| Transmitter | Not less than 250 watt with external power Booster or Not less than 250 with external 5 KW |
| Current Output option | 2.5 Amps or more |
| Interface |  |
| LCD Display | Color & day light visible |
| Interface I/O port | USB and through LAN |
| Memory Capacity | Greater than or equal to 20,000 readings |

## Item 18: Time Domain Electromagnetic (TEM) Equipment

| Feature | Specification |
| --- | --- |
| General Features |
| Operating temperature | -20° C + 65° C |
| Specific Features |
| Time gates | 48 geometrically spaced |
| Time range | Up to 16000 us or more |
| Transmitter’s current | Max. 10 A or more |
| Transmitter Loop size | 1600-2500 m2 |
| Sounding Depth Range | Up to 300m or more in favorable condition |
| Sensitivity | ≈ 0.1μV |
| Stacking | Up to 65000 stack in single loop |
| Display | Note book or Hand held PC |
| Weight(kg) | Compact and light weight |
| Transmitter Protection | Electronic and electromechanical protection. |
| Casing | The unit should be compact and light with generating and measuring block with the internal battery housed in a water proof case. |
| Accessories: |  |
|  | Cable for connection of the computer with port RS 232. |
|  | Cable for measurements with the single receiving –generating antenna. |
|  | Cable for measurements with the receiving and generating antenna (4 wires). |
|  | Cables for connection of the external battery. |
|  | Test –coil. |
|  | Charging device with a cable. |
|  | Antenna 50m X 50m (100m cable R=< 2 Ohm and 2X50 = 100 m, R= < 4 Ohm). |
|  | Additional sockets for antennas. |
|  | Dedicated Software. |

## Item	19 Resistivity Meter (Signal Averaging System)

| Feature | Specification |
| --- | --- |
| General Features |
| Power input | 12/24 V rechargeable batteries |
| Display | Alphanumeric LCD |
| Specific Features |
| Power Out | 100 watts or more |
| Current | up to 2 Amp |
| Frequency | Less than or equal to 0.8 Hz |
| Noise rejection | 95 db or more |
| Potential measuring | 10 micro volts |
| Range Resolution Range selection | Automatic |
| Resistance range | 10 micro ohms to 10 Kilo ohms |
| SP Cancellation | Automatic |
| Dynamic range | Better than 15 bits |
| Data averaging | up to 64 cycles |
| Input Impedance | 10 M ohms or more |
| Accuracy | 1% |
| Interface |
|  | User friendly menu operation with feather touch key pad |
|  | Provision for data transfer to any window based PC |
|  | Provision to display error signals in case of poor electrode contacts or discontinuity |

## Item	20 : Multi-Parameter Digital Geophysical Logger (500 m):

| Feature | Specifications |
| --- | --- |
| General module and Acquisition Console |
| Power supply | 230 Volt ± 5% by Generator driven at 50 hertzs AC frequency, suitable horse power Generator which enable to take load of 500m cable in viscous mud fluid for measuring during down and upward below 500m borehole. Generator may bear all weather temperature of Indian tropical climate condition during running condition. |
| Communication Interface | RS 232 or Latest for communication with latest Laptop |
| Software | Acquisition and interpretation Software’s with editing and presentation facility . Instantly plot the recorded graph during measurement and facility to store in external drive like Pen-drive or better storage facility. Data collected by the down hole & Up-hole probes are digitally stored during acquisition in a Laptop PC, Low resolution field printouts are produced while the data is being acquitted allowing the operator to review the data for completeness Later, appropriate scale are chosen and filters may be applied and high resolution printouts are made. Presentation quality logs from several probes are merged on the final printouts. |
| Front Panel | Voltmeter, current meter, torque, Speedo meter and Depth indicator etc will be fitted in front panel along with operating of which drive system. Or suitable device may be given for above monitoring. |
|  | Operating system of modules, which drive and borehole location should be aligned after mounting the Logger. So that operator may observe up to borehole during operation of GP Logging. |
| Operating Temperature | All weather Indian tropical condition including natural borehole temperature up to 500m (Hot water borehole to icing condition) or better. |
| Storage Temperature | -5 to 60˚C or better (All should be fitted in suitable vehicle with capacity of sitting of operators at least three persons in comfort, all equipments with all probes, tool box and Generator etc). |
| Winch Assembly |
| Motor | Suitable capacity, SCR type, motor can take load up to 500m in borehole filled with viscous mud fluid during logging operation. |
| Controller | SCR controller, with 10 amp current limit or better |
| Speed | 0 to 30 m/min |
| Maximum cable capacity | 500m |
| Emergency Brake | Winch assembly should be provided with emergency brake & provision to drive manually in case failure of Generator |
| Caliper tool |
| Number of Arms | Three or four (Two set short and long arms) |
| Diameter | ≤ 60mm |
| Length | ≤150cm |
| Measurement range | Up to 31cm approx. OR more by short arms and long arms up to 50.4cm approx. OR more |
| Probes and Tools |
| Fluid Resistivity measuring range | 0-100 ohm meter |
| Temperature measuring range | All borehole natural temperature from hot to icy condition. |
| Fluid Resistivity Resolution | ≤0.05%. |
| Temperature Resolution | ≤0.1˚C |
| Accuracy | ≤1% |
| Probe Diameter | 60 mm or lower |
| Gamma tool: | Provision for various time constants desirable |
| Electric tool: |  |
| Low range normal Resistivity range | 0-250 ohm-meter |
| High range normal Resistivity range | 0-10000 ohm-meter |
| Resistivity accuracy | ≤1% |
| Resistivity Resolution | ≤0.02% |
| Self Potential range | ≤ -1.5to 1.5 VDC |
| Self Potential accuracy | ≤1% |
| Self Potential resolution | ≤0.02% |
| Calibration Box | Calibration box of above measurement parameters |
| Mud Resistivity Meter | Portable, measured range 0 to 250 ohm meter |
| Water Analysis Tool | Portable Water analysis Tool |
| Tool Kit | Complete geophysical tool kit for field operations including tripod and digital Multi-meter |
| Spares | Acquisition system spares including complete cable head spares, probe parts, winch parts, O-rings and consumables |

## Item	21 : Multi-Parameter Digital Geophysical Logger (1000 m)

| Feature | Specifications |
| --- | --- |
| General module and Acquisition Console |
| Power supply | 230 Volt ± 5% by Generator driven at 50 hertzs AC frequency, suitable horse power Generator which enable to take load of 1000m cable in viscous mud fluid for measuring during down and upward below 1000m borehole. Generator may bear all weather temperature of Indian tropical climate condition during running condition. |
| Communication Interface | RS 232 or Latest for communication with latest Laptop |
| Software | Acquisition and interpretation Software’s with editing and presentation facility . Instantly plot the recorded graph during measurement and facility to store in external drive like Pen-drive or better storage facility. Data collected by the down hole & Up-hole probes are digitally stored during acquisition in a Laptop PC, Low resolution field printouts are produced while the data is being acquitted allowing the operator to review the data for completeness Later, appropriate scale are chosen and filters may be applied and high resolution printouts are made. Presentation quality logs from several probes are merged on the final printouts. |
| Front Panel | Voltmeter, current meter, torque, Speedo meter and Depth indicator etc will be fitted in front panel along with operating of which drive system. Or suitable device may be given for above monitoring. |
|  | Operating system of modules, which drive and borehole location should be aligned after mounting the Logger. So that operator may observe up to borehole during operation of GP Logging. |
| Operating Temperature | All weather Indian tropical condition including natural borehole temperature up to 1000m (Hot water borehole to icing condition) or better. |
| Storage Temperature | -5 to 60˚C or better (All should be fitted in suitable vehicle with capacity of sitting of operators at least three persons in comfort, all equipments with all probes, tool box and Generator etc). |
| Winch Assembly |
| Motor | Suitable capacity, SCR type, motor can take load up to 1000m in borehole filled with viscous mud fluid during logging operation. |
| Controller | SCR controller, with 10 amp current limit or better |
| Speed | 0 to 30 m/min |
| Maximum cable capacity | 1000m |
| Emergency Brake | Winch assembly should be provided with emergency brake & provision to drive manually in case failure of Generator |
| Caliper tool |
| Number of Arms | Three or four (Two set short and long arms) |
| Diameter | ≤ 60mm |
| Length | ≤150cm |
| Measurement range | Up to 31cm approx. OR more by short arms and long arms up to 50.4cm approx. OR more |
| Probes and Tools |
| Fluid Resistivity measuring range | 0-100 ohm meter |
| Temperature measuring range | All borehole natural temperature from hot to icy condition. |
| Fluid Resistivity Resolution | ≤0.05%. |
| Temperature Resolution | ≤0.1˚C |
| Accuracy | ≤1% |
| Probe Diameter | 60 mm or lower |
| Gamma tool: | Provision for various time constants desirable |
| Electric tool: |  |
| Low range normal Resistivity range | 0-250 ohm-meter |
| High range normal Resistivity range | 0-10000 ohm-meter |
| Resistivity accuracy | ≤1% |
| Resistivity Resolution | ≤0.02% |
| Self Potential range | ≤ -1.5to 1.5 VDC |
| Self Potential accuracy | ≤1% |
| Self Potential resolution | ≤0.02% |
| Calibration Box | Calibration box of above measurement parameters |
| Mud Resistivity Meter | Portable, measured range 0 to 250 ohm meter |
| Water Analysis Tool | Portable Water analysis Tool |
| Heavy duty tripod assembly | Can take load of 1000 m armored cable up to depth 1000 m in viscous mud filled fluid |
| Tool Kit | Complete geophysical tool kit for field operations including tripod and digital Multi-meter |
| Spares | Acquisition system spares including complete cable head spares, probe parts, winch parts, O-rings and consumables |

## Item	22 : Resistivity Meter (Indigenous or Equivalent)

| Feature | Specification |
| --- | --- |
| General Features |
| Input Power Source | 12/24V rechargeable batteries |
| Display with the System | Alphanumeric Liquid Crystal Display. |
| Specific Features |
| Power Output | 40 Watts or more |
| Noise Rejection | 95 db or more |
| Potential Measuring range | 10 micro volts. |
| Range selection | Manual/Automatic |
| Resistance Measuring Range | 10-3 to 104 ohms. |
| Self potential cancellation | Automatic |
| Data Averaging | Upto 16 cycles or more |
| Input Impedance | 1 Mega Ohm or more |
| Accuracy | + 1% |
| Output | Resistance |
| Protection | Protected against circuit overloads. |
| Weight and Dimension | Light weight and small |

## Item	23: Snow Water Equivalent (SWE)-Snow Pillow

| Feature | Specification |
| --- | --- |
| Site Conditions |
| Ambient Temperature | From -40 to +60 |
| Humidity | 5 to 100 % |
| Altitude | 2000 to 5500 meter |
| Snow Pillow | For measurement of snow water equivalent Consisting of liquid-filled pillow and pressure transducer ( or, alternatively, a system consisting of a standpipe, float and shaft encoder) Four snow pillow per station plumbed together Total area min 7 m2 (80 sqft) Tanks made from stainless steel by manufacture experienced with fabricating snow pillow tanks. Antifreeze solution for filling snow pillow Pipes and valve as required |
| Range | 1000 mm water equivalent |
| Pressure measuring Accuracy | 1% full scale (10mm) |
| General Features |
| Output Interface | SDI12/ RS 485 / / 4-20 mA / Compatible with Data logger |
| Power Supply | 12 V DC or switch rated for 12 VDC |
| Material | Corrosion Resistance Metal (Stainless steel or Aluminum) |
| Enclosure | NEMA 4 |
| Tools | Complete tool kit for operation and routine maintenance |
| Manuals | Full Documentation and maintenance manual in English |
| Accessaries | Sensor Mounting support, cables and other accessories as required |

Item 24: Electromagnetic Flow Meter

| Feature | Specification |
| --- | --- |
| Accuracy | +- 1 % of reading plus zero stability |
| Max Water Velocity | 10 m/s |
| Tube Material | 316 Stainless Steel |
| Electrode Material | AISI 316L (Standard) / Hastelloy/ Titanium |
| Liner Material | PTFE |
| O-Ring Seal Material | Viton and Buna N |
| Flow range | 0.75 to 63 L/s |
| Temperature Range | 0 Degree C to 80 Degree C |
| Max Pressure | 16 bar |
| max cable Length | 100 m |
| Min Conductivity | 5 microS/cm |
| Rating | IP 68 |
| CE Declaration | EN 61326:1997 to EN 61326/A3:2003 |
| Power | 12/24 V DC ; 90-264 VAC / Battery |
| Datalogger | Builtin |
| Graphic Display | With totalizer, indication above 8 mm |
| Power Management | Total management with automatic sleep function |
| Communication | MODBUS RTU on RS 485 |
| Password | Multi level |
| Line Size | 2 Inch to 6 Inch, to be specified later by Implementing Agencies |

### Item 25: Water Quality Sondes

| Feature | Value |
| --- | --- |
| Site Conditions |
| Ambient Temperature | -5 to 45 Degree C |
| Humidity | 5-100 Percentage |
| Altitude | 0-5000 meter |
| Multi parameter Sonde |
| Ports | 2 optical, 1 conductivity/temperature, 1 pH |
| Response Time | <90 s |
| Output | SDI-12, RS-232 or compatible with handheld device |
| Tools | Complete tool kit for installation and routine maintenance |
| Manuals | Full documentation and maintenance instructions in English |
| Depth Sensor |
| Accuracy | 0.003 m |
| Resolution | 0.001 m |
| Range | 0 to 60m |
| Conductivity |
| Accuracy | +/- 3% FS or 5μS/cm |
| Resolution | 1μS/cm |
| Range | 0 - 100 μS/cm |
| Dissolved oxygen |
| Sensor Type | Optical |
| Accuracy | +/- 5% reading or +/- 0.2 mg/L |
| Resolution | 0.01 mg/L |
| Range | 0 to 50 mg/L |
| Sensor Cleaning | Automated sensor cleaning mechanism |
| Temperature |
| Accuracy | +/- 0.2oC |
| Resolution | 0.2oC |
| Range | -5 to 45o C |
| Turbidity |
| Accuracy | +/- 5% reading or 2 NTU |
| Resolution | 1 NTU |
| Range | 0 to 1000 NTU |
| Sensor Cleaning | Automated sensor cleaning mechanism |
| pH |
| Accuracy | +/- 0.2 pH units; +/- 1.0 mV |
| Resolution | 0.01 pH unit; 0.1 mV |
| Range | 2 - 12 pH units (minimum) ; 0-14 pH units (Preferred) |

### 1. Equipment with Integrated features

Some manufacturers produce equipment with more than one item integrated into one equipment. The examples of such integration could be Shaft Encoder with Datalogger, Rain Gauge with Datalogger, Datalogger with GSM Modem, Datalogger with Satellite Transmitter etc.

The proposer is free to offer integrated equipment having features of more than one item. However, the integrated equipment must satisfy all the specification requirement of individual items. For example, if a proposer offers a equipment model with Data logger as integral part of Shaft encoder, the proposed equipment must meet all specification requirement of Shaft Encoder AND Datalogger.

The examples of integrated equipment provided above are indicative only; the proposer is free to offer any combination of equipment as integrated equipment.

# SECTION V-A

# COMPLIANCE MATRIX

For Each clause of Technical Specifications the Proposer shall prepare a table below:

| Feature | Technical Specifications desired | Technical Specification Offered | Compliance (yes /no) | Make of the equipment Proposed | Model No.of the equipment Proposed | Page Number in the Proposal submitted where documentary evidence is enclosed |
| --- | --- | --- | --- | --- | --- | --- |
| Item 1 Automatic Rain Gauge |
| 1. Sensor Type | Tipping Bucket type with Reed Switch | Tipping Bucket type with Reed Switch | Yes/No | XXX Company | YYY Model | 37 |
| 2. Capacity | 250 mm/hour or better | 200 mm/hour | No | XXX Company | YYY Model | 42 |
| 3. Resolution | 0.5 mm or better | 0.2 mm | Yes | XXX Company | YYY Model | 47 |
| 4. |  |  |  |  |  |  |
| 5. |  |  |  |  |  |  |
| 6. |  |  |  |  |  |  |

# APPENDIX 1: Details of Existing Earth Receiving Stations and ISRO Satellite System

## Features of ISRO TDMA transmission

Features of ISRO TDMA transmission scheme are provided for general guidance.  However international norms applicable for TDMA may be followed.

- Total number of DCU that could be accommodated in a single carrier is 1800.

- By including CRC in the data frame, data validity could be ensured.

- With preserving BCH coding of SID, data quality could be checked and valid data retrieved even for the bad CRC.

- By preserving present SID (Station Identification Code) structure of IMD, SID for all users of DRT could be standardized.  The SID consists of 21bits (9 bits for user type, 2 bits for priority, and 10 bits for Platform ID).

- With Forward error correction convolution coding, better data quality is ensured.

- With one repeat transmission, reliability of data reception is improved.

| 1 | CRC CODE GENERATION | Polynomial; CRC-CCITT-16 X16+X12+X5+1 |
| --- | --- | --- |
| 2 | DATA SCRAMBLING | Polynomial: 1+X-1+X-15 Initial State: 6959 (Hex) |
| 3 | CONVOLUTION ENCODING | Convolution Coding ½ Rate, Constraint Length K=7 Polynomial: G1=133(Octal), G2=171(Octal) |
| 4 | HEADER DETAILS | CR: 192 Symbols (all ’0’s) BTR: 64 Symbols (all ’1’s) UW: 64 Symbols (07EA CDDA 4E2F 28C2 (Hex)) Note: UW transmitted with LSB first of every byte, starting from 07EA. (See Fig.1) |
| 5 | RF DATA ENCODING | Differential coding (NRZ-L) is done for the entire burst (Preamble and the convolution coded bits) before RF modulation. |

Fig 1: Burst Transmission Format for TDMA Technique  (4800 Symbols/sec.)

Fig 2: TDMA Transmission Frame Format

Fig.1 may be referred to. CRC is calculated for 262 bits which include FS and EOT. It is then scrambled. 1byte, all ’0’s is added with the scrambled bits, after which the entire bits are convolution coded. Preamble (CR, BTR and UW) is appended with the convolution coded bits. The resulting bits are then differential coded and transmitted.

The system should have flexibility to accommodate more number of carrier channels by suitable changes in the TDMA transmission scheme.

Additional details, if required, will be provided at the time of the design review meeting which will be held with the successful tenderer.  However, it should be ensured by the tenderer that the system configuration is flexible and accommodate more than 30 sensors without any additional cost.

## INSAT DRT Specifications

For the purpose of data transfer from field DCU to Data Receiving Earth Station at New Delhi, the Data Relay Transponder (DRT) on the different INSAT/ KALPANA-1 series of satellites shall be used and the specifications given below shall be treated as standard  to be adhered by the offered telemetry system.

| SATELLITE | KALPANA-1 (740 E) | INSAT-3A (830E) | INSAT-3D (820E) |
| --- | --- | --- | --- |
| RECEIVE FREQ. BAND | 402.65 - 402.85Mhz | 402.65 - 402.85Mhz | 402.10 - 402.50Mhz |
| TRANSMIT FREQ. BAND | 4500-4510Mhz band 4506.05Mhz | 4500-4510Mhz band 4506.05Mhz | 4500-4510Mhz band 4506.05Mhz |
| RECEIVE G/T | -19db/deg.K | -19db/deg.K | -19db/deg.K |
| MAX.EIRP | 24dBW peak | 24 dbW peak | 24 dbW peak |
| C-BAND EIRP for RECEIVE FLUX DENSITY | 2.0dBW for –146 dBW/m2 | 2.0 dbW for –146 dbW/m2 | 2.0 dbW for –146 dbW/m2 |
| REC.POLARISATION | RHCP | LHCP | LHCP |
| TRANSMIT POL | LINEAR | LINEAR | LINEAR |
| FREQ.TRANSLATION ERROR | ( 40Khz over life ( 6Khz over 1 month | ( 40Khz over life ( 6Khz over 1 month | ( 40Khz over life ( 6Khz over 1 month |

Data Relay Transponder (DRT) onboard INSAT 3D will have a receiving frequency band of 402.3  MHz ( 200 KHz.

## Channel Specifications for TDMA transmission format

Table below gives the present AWS parameters and their identification code used in the TDMA transmission format.

| Sl.No. | Channel No. | Identification Code | Parameter |
| --- | --- | --- | --- |
|  | 1 | 0000 (:0) | Instantaneous sampled value of air temperature in deg C at the end of every full hour UTC. |
|  | 2. | 0001(:1) | Water level sampled at end of every full hour IST |
|  | 4 | 0100(:4) | Wind speed in knots (3 minute vector averaging prior to full hour UTC). |
|  | 5 | 0101(:5) | Wind direction in degrees (3 minute vector averaging prior to full hour UTC). |
|  | 7 | 0111(:7) | Instantaneous value of RH at the end of every full hour UTC. |
|  | 10 | 1110(:14) | Duration of bright sunshine since last 20 UTC. Reset to zero at 20 UTC. (Global radiation will be transmitted in this slot instead of duration of sunshine. |
|  | Cal1 | :C1 | Battery voltage (volts) |
|  | Cal2 | :C2 | Hourly rainfall (rounded off to next higher integer). |

# SPECIFICATIONS OF EXISTING EARTH RECEIVING STATION EQUIPMENTS

## Details of Earth Receiving Station Antenna

|  | Reflector size | : | 3.8 metres |
| --- | --- | --- | --- |
|  | Reflector type | : | Solid fibre glass material |
|  | Mount Design | : | Polar mount/ any other suitable design (TBS) |
|  | Feed Mount | : | Prime focus feed |
|  | Feed type | : | Linear |
|  | Input frequency (for feed) |  | 4.5 to 4.8 GHz |
|  | G/T | : | 31.7 dB /(K |
|  | Operating frequency | : | 4500- 4800 MHz |
|  | Gain | : | >43 dB |
|  | Polarization | : | LHCP / RHCP selectable |
|  | Elevation Adjustment Range | : | 0-90(( Coarse & fine adjustment) Angles to be engraved on the antenna |
|  | Azimuth Adjustment Range | : | 0-360((Coarse & fine adjustment) Angles to be engraved on the antenna |
|  | Wind loading : a) Operational | : | 100 KMPH or better |
|  | b) Survival | : | 175 KMPH or better |
|  | Operating rainfall rate | : | 100mm/hr and water proof. |

## Features of LNA

|  | Frequency range | : | 4500 – 4800 MHz |
| --- | --- | --- | --- |
|  | Bandwidth | : | 300 MHz(typical) |
|  | Noise temperature (Ambient Temp. 25°C) | : | ≤50 (K (45 (K typical) |
|  | Gain | : | > 60 dB |
|  | Gain ripple | : | Not more than (0.5 dB (Over entire 300 MHz pass band) |
|  | Max. RF input | : | -50 dBm composite |
|  | Max. RF input with no damage | : | 0 dBm CW in pass band |
|  | Input / Output VSWR : (4.5 GHz to 4.8 GHz ) | : | 1.2 : 1 |
|  | Operating Temperature | : | -10 to 55 oC |
|  | Humidity | : | 0-100 per cent with condensation |

## Features of Synthesized Down Converter

The general features are listed below.

RF input	:	4500 – 4800 MHz

IF output range	:	Compatible to Demodulator Input

(May be 100-180 MHz)

RF input level	:	-55 dBm typical

IF output level	:	+20 dBm at 1 dB compression

Frequency stability over time	:	+/- 1 x 10-9 / day

Frequency stability over temperature	:	+/- 1 x 10-8 / day

# Section VI.  Sample Forms

## 1.  Proposal Form (for first stage competition)

Date:	[ insert:  date of Proposal ]

Loan/Credit No.:	[ Purchaser insert: number ]

[ Purchaser to specify:  “IFP No.:	[ number ]” ]

[ insert:  name of Empanelment ]

To:  [ Purchaser insert:  Name and address of Purchaser ]

Dear Sir or Madam:

Having examined the Document for setting up Empanelment, including Addenda Nos. [ insert numbers ], the receipt of which is hereby acknowledged, we, the undersigned, offer to supply and deliver the Goods under the above-named Agreement in full conformity with the said document for setting up Empanelment for the sums as may be determined in accordance with the terms and conditions of the Purchase Orders to be issued under the above mentioned Empanelment.

We undertake, if our Proposal is accepted, to deliver the Goods in accordance with the tentative delivery schedule and list of consignee indicated in the Schedule of Requirements (exact requirements to be further specified in the Purchase Orders). The Goods are sourced from <name of the country>.

If our Proposal is accepted, we undertake to provide a performance security in the form, in the amounts, and within the times specified in the document for setting up Empanelment.

We agree to abide by this Proposal, for the Proposal Validity Period specified in the document for setting up Empanelment and it shall remain binding upon us and may be accepted by you at any time before the expiration of that period.

Until the formal final Empanelment is prepared and executed between us, this Proposal, together with your written acceptance of the Proposal and your notification of award for setting up the Empanelment, shall constitute a binding Contract between us.  We understand that you are not bound to accept any Proposal you may receive.

We undertake that, in competing for (and, if the award is made to us, in executing) the above agreement, we will strictly observe the laws against fraud and corruption in force in India namely “Prevention of Corruption Act 1988”.

We hereby certify that we have taken steps to ensure that no person acting for us or on our behalf will engage in bribery.

Commissions or gratuities, if any, paid or to be paid by us to agents relating to this Proposal, and to agreement execution if we are awarded the Agreement, are listed below:

| Name and Address of Agent |  | Amount in Indian Rupees |  | Purpose of Commission or Gratuity |
| --- | --- | --- | --- | --- |
| (if none, state “none”) |

We confirm that we comply with the eligibility requirements as per ITP clause 4 of the document for setting up Empanelment.

Dated this [ insert: number ] day of  [ insert: month ], [ insert: year ].

Signed:

Date:

In the capacity of [ insert:  title or position ]

Duly authorized to sign this Proposal for and on behalf of [ insert:  name of Proposer ]

Checklist for documents to be submitted with first stage Proposal form:

| Document | Attached (Yes / No) |
| --- | --- |
| Power of Attorney in favour of the signatory of Proposal |  |
| Proposal Security in the required amount and format |  |
| Certificate of incorporation of the Proposer |  |
| Performance statement as per required format, along with copies of purchase order, client certificate etc. |  |
| Copies of balance sheet & Profit & Loss statement for last 3 fiscal years |  |
| No deviation statement on technical specification |  |
| Declaration of agreement with all terms and condition of the document for setting up Empanelment |  |
| Manufacturer’s authorization (if item is not manufactured by the Proposer) |  |
| Any other documents as per ITP 7.1 and 14 |  |

## 2 - Form of Bank Guarantee

Proposal Security Form/ Bid Security (Bank Guarantee)

________________________________
[insert: Bank’s Name, and Address of Issuing Branch or Office]

Beneficiary: Dy Director ,River Data Directorate , Central Water Commission, Wing -4, 2nd Floor, West Block – I, R.K.Puram, New Delhi – 110605., Tel: +911126100285,+911126108075 Fax: +911126181267,  Email: rdcdte-cwc@nic.in

Date:[insert: date]

BID GUARANTEE No.:[insert: Bid Guarantee Number]

We have been informed that [insert: name of the Bidder] (hereinafter called "the Bidder") has submitted to you its bid dated [insert: bid date] (hereinafter called "the Bid") for the execution of [insert: name of contract] under Invitation for Bids No. [insert: IFB number].

Furthermore, we understand that, according to your conditions, bids must be supported by a bid guarantee, and that the bid guarantee automatically covers any alternative bids included in the Bid, if the Bidder is permitted to offer alternatives and does so.

At the request of the Bidder, we [insert: name of Bank] hereby irrevocably undertake to pay you any sum or sums not exceeding in total an amount of [insert: amount in figures]([insert: amount in words]) upon receipt by us of your first demand in writing accompanied by a written statement stating that the Bidder is in breach of its obligation(s) under the bid conditions, because the Bidder:

(a)	has withdrawn the Bid (or any parts of it) during the period of bid validity specified by the Bidder in the Bid Submission Form or any extension of the period of bid validity which the Bidder subsequently agreed to; or

(b)	having been notified of the acceptance of the Bid by you during the period of bid validity, (i) failed or refused to execute the Contract Agreement, or (ii) failed or refused to furnish the performance security, if required, in accordance with the Instructions to Bidders.

This guarantee will expire: (a) if the Bidder is the successful bidder, upon our receipt of copies of the contract signed by the Bidder and the performance security issued to you upon the instruction of the Bidder; or (b) if the Bidder is not the successful bidder, upon the earlier of (i) our receipt of a copy of your notification to the Bidder of the name of the successful bidder; or (ii) twenty-eight days after the expiration of the Bid's validity.

Consequently, any demand for payment under this guarantee must be received by us at the office on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees,ICC Publication No. 458.

_____________________________

[Signature(s)]

[Note to Bidders:  Instructions on amount and currency can be found in the ITB Clause and BDS for "Securing the Bid."  Joint Ventures need to also ensure that their Bank Guarantee meets the requirements for Joint Ventures as provided in the same Clause.]

## 3.  Manufacturer’s Authorization Form

(Manufacturer’s or Producer’s letterhead)

To:  [ insert:  name of the Purchaser ]

WHEREAS [ name of the manufacturer or producer ] (hereinafter, “we” or “us”) who are established and reputable manufacturers or producers of [ name and/or description of the Goods requiring this authorization ] (hereinafter, “Goods”) having production facilities at [ insert:  address of factory ]  do hereby authorize [ name and address of Proposer ] (hereinafter, the “Proposer”) to submit a Proposal, and sign the Agreement with you against IFP [ title and reference number of the Invitation for Proposals ] including the above Goods produced by us.

We hereby extend our full guarantee and warranty for the above specified Goods against this document for setting up Empanelment.

For and on behalf of the Manufacturer or Producer

Signed: ___________________________________________________________

Date: ______________________________________

In the capacity of [title, position, or other appropriate designation] and duly authorize to sign this Authorization on behalf of [name of manufacturer or producer]

Note:	This letter of authority should be on the letter head of the manufacturers and should be signed by a person competent and having the power of attorney to legally bind the manufacturer.  This should be included by the Proposer in its Proposal.

|  |  | 4. Proforma for performance statement (For a period of last five years) |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| IFP No: ___________ | Date of Opening: ______________ | Time :_____________Hours |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  | Name of the Firm :____________________________________________________ |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Order Placed By | Order No. | Description and quantity | Value of order | Date of completion of delivery | Remarks indicating | Was the supply of goods |  |  |  |  |  |  |  |  |  |  |  |  |  |
| (Full address of | and Date | of ordered |  | As per contract | Actual | reasons for late |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Purchaser) |  |  |  |  |  | delivery, if any | satisfactory ? |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  | (Attach a certificate from |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  |  |  |  |  | the Purchaser/Consignee) |  |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  | Signature and seal of the Proposer |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

# Footnotes

Any questions regarding this list should be addressed to the Director, Procurement Policy and Services Group, Operational Core Services Network, the World Bank.

# Text boxes

1 hour frame with slot size 1 second (One hour frame is divided into 6 time windows, each of 10 minute duration)

1 Hour