## Slide 1: Water Quality-A NHP Perspective
Rajendra Mohan Bhardwaj
Water Quality Expert
National Hydrology Project
rmbhardwaj@gmail.com

## Slide 2
Overview of Presentation
Requirement of Water Quality Monitoring
Planning and Execution of Water Quality Monitoring
Guideline for setting up of Water Quality Laboratory
(Levels-I to III/II+ Laboratories)

## Slide 3: Requirement of Water Quality Monitoring

## Slide 4
Requirement of Water Quality Monitoring
  Understand physical, chemical and biological properties
  Managing freshwater ecosystem
  Compliance to guidelines or targets
  How will Information be used

## Slide 5: Water Use and Properties
Properties of the water govern its suitability for different uses
“Uses” include: drinking water, irrigation, recreation, mass bathing, industrial, aquatic ecosystem, assimilating wastewaters etc.
Each use has its own set of water quality requirements

## Slide 6
Water Quality Criteria
Primary water Quality Criteria-1979 (Ambient)
Tolerance Limits for Industrial Effluents-BIS-2490-1981
Bathing Water Quality Criteria-2000 (Ambient)-EPA-1986
Irrigation Water Quality Criteria-BIS-11624-1986
Coastal Water Quality Criteria-2000-EPA 1986
Drinking Water Quality Criteria-BIS-10500-1991 & 2012
Industrial water use criteria-BIS-201,2724,4251,6582
General Discharge Standards-1986-EPA,1986
Industry Specific Discharge Standards, EPA,1986
WHO-Guidelines for Drinking Water-1958 onwards-2019

## Slide 7: Choosing What to Monitor
Basic parameters of water quality
  pH, Temp., Suspended Solids, Turbidity, Conductivity (or TDS), Major Ions
Demand Parameters
  DO, BOD, COD
Pathogens
  Coliform – Faecal & Straptococci
Nutrients
  Nitrate, Nitrite and Phosphates
Toxic Contaminants
  Trace metals and Pesticides

## Slide 8
Classifying Water Quality: Guidelines, Standards and Targets
Targets can be:
precise values/standards
a range of values, or
a water quality classification
Water quality classifications can be used for chemical characteristics
Complying, Non-Complying, Satisfactory, Not Satisfactory
High, Good, Moderate, Poor, Bad
Excellent, Good, Fair, Marginal, Poor
Classifications can be used to indicate whether water is suitable for particular uses

## Slide 9: Planning and Execution of Water Quality Monitoring

## Slide 10
Protocol on Water Quality Monitoring
Water Quality Monitoring is performed in accordance with the “Uniform Protocol on Water Quality Monitoring Notification” 2005, subsequently updated during 2017 as a guideline

## Slide 11
Classification of Water Quality Monitoring Stations
Water quality monitoring stations are classified as follows:
  Baseline stations - no influence of human activities on water quality
  Trend stations-particular point on a watercourse varies over time due to the influence of man made activities
  Flux stations or Impact stations - location for measuring the mass of particular parameter/pollutant/extent of pollution due to human interference

## Slide 12
Steps Involved in Planning Water Quality Monitoring
Network Design
  Selection of sampling locations
  Optimum number of locations
  Parameters to be measured
  Frequency of sampling
  Component to be samples – water, sediment or biota
Sampling
  Representative sampling
  Field testing
  Sample preservation and transport

## Slide 13
Steps Involved in Planning Water Quality Monitoring (contd.)
Laboratory Work
  Laboratory procedures
  Physical, chemical, bacteriological analysis
Data Management
  Storage
  Statistical analysis
  Presentation
  Interpretation
  Reporting
  Quality Assurance
  Production of reliable data
  Quality control- Internal AQC & External AQC

## Slide 14
Protocol for Sampling
Protocol on sampling are provided in
Bureau of Indian Standards i.e., IS-3025
Standard Methods for the Examination of Water & Wastewater- 22nd Edition [Published jointly by American Public Health Association (APHA)
American Society for Testing and Materials (ASTM)

## Slide 15
Sample Containers and Preservation
 | Analysis | Container | Volume (mL) | Preservation
1 | General (SS, TDS, major ions, chlorophyll-a) | Glass, PE | 1000 | 4oC
2 | COD, NH3, NO2-+NO3- | Glass, PE | 500 | H2SO4, pH <2
3 | P | Glass | 100 | -
4 | DO | special BOD bottle | 300 | DO fixing
5 | BOD | Glass, PE | 1000 | 4oC, Dark
6 | Coliforms | Glass, PE, Sterilised | 300 | 4oC, Dark
7 | Heavy metals (Cd, Zn) | Glass, PE | 500 | HNO3, pH <2
8 | Mercury | Glass | 1000 | HNO3, pH <2
9 | Pesticides | Glass, Teflon | 1000 | 4oC, Dark

## Slide 16: Guideline for Setting up of Water Quality Laboratory (Levels-I to III/II+ Laboratories)

## Slide 17
Mandates and Objectives of Water Quality Monitoring
Mandate
  Assessment of water resources-Quantitative & Qualitative
  Setting up of laboratories.
Objectives
  Establishment of baseline water quality
  Assessment of suitability of water for various uses
  Detection of trends in water quality changes.
  Dissemination of water quality information

## Slide 18
Concept of Development of Laboratories
Analytical laboratory are comprised of following units
  Physical
  Chemical
  Instrumentation
  Bacteriological
  Preparation Room-Balance & Pure water system
  Sample room-sample receipt & storage
  Chemical Store Room
  Staff Room

## Slide 19
Water Quality Laboratory (Levels-I to III/II+ Laboratories)
Three tier laboratory system analyse following parameters
Level-I Laboratories: Six parameters
Level-II Laboratories: 25 physico-chemical and bacteriological parameters
Level-III/II+ Laboratories: 41 parameters including heavy metals / toxic parameters and pesticides

## Slide 20
Level-I Laboratory-Parameters (6)
Temperature
Colour
Odour
Electrical Conductivity
Dissolved Oxygen
pH

## Slide 21
Level-II Laboratory-Parameters (25)
(in Addition to Level -I Parameters) | 
Carbonate | Sodium
Bicarbonate | Potassium
Chloride | Boron
Fluoride | Calcium (as Ca)
Sulphate | Magnesium (as Mg)
Nitrate | Total Dissolved Solids
Nitrite | Biochemical Oxygen Demand
Silicon | Chemical Oxygen Demand
Iron | Total coliforms
Ortho-phosphate | Faecal Coliform
 | Turbidity

## Slide 22
Level-III/II+ Laboratory-Parameters (41)
(in Addition to Level -I and Level II Parameters) | 
Zinc | BHC (α, β, γ)
Lead | DDT (OP,PP)
Chromium | Alpha Endosulphan
Mercury | Aldrin
Cadmium | Dieldrin
Nickel | Carbamate
Arsenic | Malathian
Ammonical Nitrogen | Methyl Parathian

## Slide 23
Instrumentation in water laboratory
General Instruments | 
Thermometer | BOD Incubator PSD controller
Electronic Balance | COD Digestor
Electronic Pipette Controller | Bacteriological incubator
Magnetic Stirrer | Laminar Air Flow
Refrigerator | Fume Hood
Dispensor | Auto Clave
Ultra Water Purification System/RO System | Water Bath
pH Meter | Air Conditioner
Conductivity Meter | Water Geyser (Heater)
Turbidity Meter | Digital Burette
Hot Air Oven | 

## Slide 24
Instrumentation in water laboratory (contd.)
Advanced Instruments
Ion Meter
DO Meter
Flame Photometer
Colorimeter
UV Visible Spectrophotometer
Microwave Digestor
Gas chromatography–Mass Spectrometry (GC-MS)
Atomic Absorption Spectrophotometer (AAS)
Inductively coupled plasma mass spectrometry (ICP-MS)

## Slide 25
Instruments Requirement in Level –I Laboratories
Parameter | Instrument Required
Temperature | Digital Thermometer
Colour | Nessler tubes, color comparator, Spectrophotometer
Odour | Glass stoppered bottles
Electrical Conductivity | Conductivity Meter
Dissolved Oxygen | DO Meter / Burette
pH | pH Meter

## Slide 26
Instruments Requirement in Level –II Laboratories
Parameter | Instrument Required
Carbonate Bicarbonate Chloride | Digital Burette
Fluoride | Ion Meter, Spectrophotometer
Sulphate | Ion Meter
Nitrate (N) & Nitrite (N) | Ion Meter, UV Visible Spectrophotometer
Silica | UV Visible Spectrophotometer, ICP-AES
Iron | Inductively Coupled Plasma-Atomic Emission Spectroscopy
Phosphate | UV Visible Spectrophotometer

## Slide 27
Instruments Requirement in Level –III/II+ Laboratories
Parameter | Instrument Required
Ammonia (N) | UV Visible Spectrophotometer
Zinc, Lead, Chromium, Mercury, Cadmium, Nickel, Arsenic | Atomic Absorption Spectrophotometer (AAS) / Inductively coupled plasma mass spectrometry (ICP-MS)
BHC (α, β, γ), DDT (OP,PP), Alpha Endosulphan, Aldrin, Dieldrin, Carbamate, Malathian, Methyl Parathian | Gas chromatography–Mass Spectrometry (GC-MS), GC,HPLC, Soxhlet Extraction Unit

## Slide 28
Laboratory Design-Required Units
Description | Specific requirements
Wet chemical Instrumentation Bacteriological Weighing and extractions Rinsing room Sample storage room Chemical storage room | acid proof Dust Free clean air conditioned Dust Free clean solid floor, Dust Free clean Open windows water proof cool, dry, dark

## Slide 29
Space Requirement-Typical Laboratory
Level | Area | Dimension
Level I field laboratory Level II laboratory Level III/II+ laboratory | 30m2 200m2 300m2 | 4 x 7m 13 x 16m 15 x 20m

## Slide 30
Cost of Laboratory Establishment
Level-I Laboratory | 
Particulars | Cost (in Rs.)
Instruments | 3,00,000
Glassware | 25,000
Chemicals | 20,000
Airconditioning | 50,000
Modular Furniture | 1,00,000
Miscellaneous | 5,000
Total | 5,00,000

## Slide 31
Cost of Laboratory Establishment
Level-II Laboratory | 
Particulars | Cost (in Rs.)
Instruments | 87,00,000
Glassware | 2,00,000
Chemicals | 2,00,000
Airconditioning | 1,00,000
Modular Furniture | 2,00,000
Miscellaneous | 1,00,000
Sub Total | 95,00,000

## Slide 32
Cost of Laboratory Establishment
Level-III/II+ Laboratory | 
Particulars | Cost (in Rs.)
Instruments | 1,75,00,000
Glassware | 50,000
Chemicals | 2,00,000
Airconditioning | 1,00,000
Modular Furniture | 1,25,000
Miscellaneous | 25,000
Sub Total | 1,80,00,000

## Slide 33
Manpower Requirement for Laboratory
Level-I Laboratory | 
Laboratory Incharge/ Scientist/ Research Officer | One
Chemist | One
Laboratory Assistant | One
Lab Boy/Attendant | One

## Slide 34
Manpower Requirement for Laboratory
Level-II Laboratory | 
Laboratory Incharge/ Scientist/ Research Officer | One
Chemist/Scientific Assistant/Research Assistant | Three
Microbiologist | One
Laboratory Assistant | Three
Lab Boy/Attendant | Two

## Slide 35
Manpower Requirement for Laboratory
Level-III/II+ Laboratory | 
Laboratory Incharge/ Scientist/ Research Officer | One
Chemist/Scientific Assistant/Research Assistant | Five
Microbiologist | One
Laboratory Assistant | Five
Lab Boy/Attendant | Two

## Slide 36
Requirement of Glassware
Quality of glassware is important in precise measurement
Glassware is typically selected for a particular analysis
Metrological grade glassware are recommended for high precision volumetric measurements
Metrological grade is determined by the traceability of the calibration to an NIST standard
Borosilicate/Quartz glass are preferred as they can withstand very high temperatures
Darkened brown or amber (actinic) glass are mandatory for chemicals to prohibit photochemical reaction by UV/IR

## Slide 37
Requirement of Chemicals
Quality of chemicals and reagents play a critical role in the testing of water samples
Procure the chemicals that meet all regulatory and compliance standards for their intended use
Broadly two types of chemicals/reagents used
Analytical/General (AR/GR) or laboratory reagent (LR) grade
AR/GR reagents are high purity chemicals and should only be procured for analysis of water samples rather than LR Grade

## Slide 38
Analytical Quality Control
Basic objective of a water analysis laboratory is to produce accurate/reliable analytical data through quality assurance
Essential component of a quality assurance programme is analytical quality control which refers to the routine application of procedures for controlling the measurement process.
Internal quality control/statistical quality control is the most important component of any laboratory AQC.
External quality control is applied after incorporating internal quality control practices in the laboratory and consists of the periodic analysis of reference samples.

## Slide 39
Thank You
For Your Attention