## Slide 1: Water Quality Lab, Instruments and their Specifications
By.
Dr. Raja Ram Purohit
Sc-C, MoWR, RD & GR

## Slide 2: Classification of water quality monitoring station
In accordance with the definition given in the “Uniform Protocol on Water Quality Monitoring Notification (UPWQM)” issued by the Water Quality Assessment Authority, Ministry of Environment & Forests dated 17th June, 2005 water quality monitoring stations are classified as follows:
Baseline Station: A Baseline Station means the monitoring location where there is no influence of human activities on water quality;
Trend Station: A Trend station means the monitoring location designed to show how a particular point on a watercourse varies over time due, normally, to the influence of man’s activities; and
Flux Station or Impact Station: A Flux station or Impact station means the location for measuring the mass of particular pollutant on main river stem for measuring the extent of pollution due to human interference or geological feature at any point of time and is necessary for measuring impact of pollution control measures adopted;

## Slide 3: Types of Water Quality Lab

## Slide 4: Three Tier Laboratory System
Level - I Laboratories: These Labs are the field laboratories which are located at field water quality monitoring stations on various rivers of India where in-situ values of five (05) physical parameters and one biological parameter (Dissolved Oxygen) are monitored.
Level - II Laboratories: These Labs are normally located at division offices to analyse 25 physico-chemical and bacteriological parameters.
Level - III/II+ Laboratories: These Labs are normally located in the regional level for the analysis of 41 parameters including heavy metals / toxic parameters and pesticides.

## Slide 5: A Comparison & Requirements…..
 | Level – I (6) | Level – II (25) | Level – II+ /III (41) | Analytical Instrument Required
Physical Parameters | Temp, Colour, Odour, EC/TDS, pH | Temp, EC/TDS, pH | Temp, EC/TDS, pH | Digital thermometer, pH meter, Conductivity / TDS meter, Oven, Nepheloturbidit meter
Chemical Parameters (Cations) |  | Carbonate, Bicarbonate, Chloride, Fluoride, Nitrate, Nitrite, Sulphate, Phosphate | Similar to Level-II Lab + Ammonia (N), Cynide, Total Kjeldhal Nitrogen | Digital Burette, Ion Meter & Visible Spectrophotometer
Chemical Parameters (Anions) |  | Sodium, Calcium, Magnesium, Potassium, Sillicate, Boron | Similar to Level-II Lab | Digital Burette, Flame Photometer & Visible Spectrophotometer, Water bath
Biological Parameters (Oxygen Demand) | D.O. | D.O., B.O.D & C.O.D | Similar to Level-II Lab In addition Total Organic Carbon | DO Meter / Burette, BOD Incubator & COD Digester & Total carbon Analyser
Microbiological |  | Total Colliform, Total Plate Count, F. Colliform, E. Colliform |  | Test tube with plastic cap, Bacteriological Incubator, Water Bath
Trace/Toxic Metals |  | Iron | In addition, Arsenic, Cadmium, Mercury, Chromium, Lead, Zinc, Nickel | Atomic Absorption Spectrophotometer (AAS) / Inductively coupled plasma mass spectrometry (ICP-MS)
Pesticides |  |  | Chloro-hydrocarbons (6nos) Alpha BHC, Beta BHC, Gama BHC (Lindane), OP-DDT, PP-DDT, Alpha Endosulphan, Beta Endosulpham, Aldrin, Dieldrin, 2, 4-D, Carbaryl (Carbamate), Malathian, Methyl Parathian, Anilphos | Gas chromatography–Mass Spectrometry (GC-MS)

## Slide 6: COST ANALYSIS FOR SETTING-Up of a Water Quality Lab

## Slide 7: Tentative Costs of instruments & Setting-Up of a Lab (Except Civil n Furnishing Work)
Name of the Instrument | Quantity required for each lab | Tentative cost for each in Rs. | Requirement for a Level II Lab | Requirement for a Level II+ / III lab | Cost for Level II Lab | Cost for Level II+/III Lab
Balance, Analytical (Electronic) | 2 | 50 thousand | Yes | Yes | 1,00,000 | 1,00,000
Electronic Pipette Controller | 3 | 20 thousand | Yes | Yes | 60,000 | 60,000
Magnetic Stirrer | 2 | 20 thousand | Yes | Yes | 40,000 | 40,000
Refrigerator | 1 | 30 thousand | Yes | Yes | 30,000 | 30,000
Dispensor | 5 | 20 thousand | Yes | Yes | 1,00,000 | 1,00,000
Ultra Water Purification System | 1 | 6 lakhs | Yes | Yes | 6,00,000 | 6,00,000
pH Meter | 2 | 70 thousand | Yes | Yes | 1,40,000 | 1,40,000
Conductivity Meter | 1 | 50 thousand | Yes | Yes | 50,000 | 50,000
Turbidity Meter | 1 | 1.5 lakhs | Yes | Yes | 1,50,000 | 1,50,000
Hot Air Oven | 2 | 65 thousand | Yes | Yes | 1,30,000 | 1,30,000
Digital Burette | 4 | 70 thousand | Yes | Yes | 2,80,000 | 2,80,000
Ion Meter | 1 | 1.5 lakhs | Yes | Yes | 1,50,000 | 1,50,000
UV Visible Spectrophotometer | 1 | 5-8 Lakhs | Yes | Yes | 6,00,000 | 6,00,000
Flame Photometer microprocessor based with PC controlled | 1 | 85 thousand | Yes | Yes | 85,000 | 85,000
DO Meter | 1 | 80 thousand | Yes | Yes | 80,000 | 80,000
BOD Incubator PSD controller | 1 | 90 thousand | Yes | Yes | 90,000 | 90,000
Colorimeter | 1 | 1.2 lakhs | Yes | Yes | 1,20,000 | 1,20,000
COD Digistor | 1 | 1.5 lakhs | Yes | Yes | 1,50,000 | 1,50,000
Bacteriological incubator | 1 | 80 thousand | Yes | Yes | 80,000 | 80,000

## Slide 8: Tentative Costs of instruments & Setting-Up of a Lab (Except Civil n Furnishing Work) continued….
Name of the Instrument | Quantity required for each lab | Tentative cost for each in Rs. | Requirement for a Level II Lab | Requirement for a Level II+ / III lab | Cost for Level II Lab | Cost for Level II+/III Lab
Laminar Air Flow | 1 | 50 thousand | Yes | Yes | 50,000 | 50,000
Fume Hood in SS | 1 | 80 thousand | Yes | Yes | 80,000 | 80,000
Auto Clave | 2 | 45 thousand | Yes | Yes | 90,000 | 90,000
Water Bath PSD controller | 2 | 35 thousand | Yes | Yes | 70,000 | 70,000
Ion Chromatograph | 1 | 40 Lakhs | No | Yes | 0 | 40,00,000
Atomic Absorption Spectrophotometer (AAS) with GTA and VGA | 1 | 45 lakhs | No | Yes | 0 | 45,00,000
Total Organic Carbon Analyser (TOC) | 1 | 42 Lakhs | No | Yes | 0 | 42,00,000
Inductively coupled plasma mass spectrometry (ICP-MS) | 1 | 1.20 Crores | No | Yes | 0 | 1,20,00,000
Gas chromatography–Mass Spectrometry (GC-MS) with Head Space and Purge & Trap with Auto sampler | 1 | 80 Lakhs | No | Yes | 0 | 80,00,000
Air Conditioner | 4 | 35 thousand | Yes | Yes | 1,40,000 | 140,000
Water Geyser (Heater) | 1 | 8 thousand | Yes | Yes | 8,000 | 8,000
Software (Aqua Chem etc) | 1 | 1 Lakh | Yes | Yes | 1,00,000 | 1,00,000
Tentative Cost for Lab Instrumentation |  |  |  |  | 35,73,000 | 3,62,73,000
In addition, you may require cost of Man Power and Glasswares and consumables..

## Slide 9: GENERAL SPECIFICATION OF THE ANALYTICAL INSTUMENTS

## Slide 10: pH Meter
Bench top pH meter with pH, mV meter for routine application microprocessor based,
Mains / Battery or Dual power supply (AC: 220V ±10%; 50Hz.)
pH combined electrode with milli Volt and temperature probe,
Temperature range 0 to 600C;
pH range between –2.000 to 16.000
Resolution: 0.001pH,
Accuracy : ± 0.001pH;
pH calibration up to 3 point Calibration with 2 custom and 7 standard buffers;
Beakers, stand and other standard accessories.

## Slide 11: Conductivity Meter (EC)
Bench top conductivity meter with for routine application, microprocessor based.
Mains / Battery or Dual power supply (AC: 220V ±10%; 50Hz.)
Conductivity Cell : 1.00 & 0.5 cell constant
Electrode with temperature probe, temperature range 0 to 800C;
Conductivity range: 0 to 19.99, 199.9, 1999; 0 to19.99 micro Siemens /cm and 0 to199.9 milli Siemens /cm,
Accuracy : ± 1% full scale,
Temperature compensation :0 to 80 0C .
Display: large LED
99 data memory with real time clock and data
Stand and other standard accessories.

## Slide 12: Dissolved OXYGEN KIT
Display : LCD
Range: D.O.: 0 to 20.0 mg/l and Temp: 0 to 500C
Accuracy: Dissolved Oxygen: ±0.4 mg/L, Oxygen in Air: ±0.7% O2 (for reference only), Temperature: ±0.8 ºC/1.5 ºF
Power Off: Auto shut-off
Data Hold: Holds the current displayed Reading
Operating Temp: 0 to 50 ºC, 32 to 122 ºF
Sample Time: Approx. One (1) Second
Operating Humidity: Less Than 80% RH
Power Supply: 4-AAA Alkaline Batteries

## Slide 13: DIGITAL BURETTE & PIPETTE
Burette
  Recirculation valve for priming without loss of reagent
  Dispensing range from 0.01 to 999.9 ml
  Simple calibration program Discharge tube rotation 3600
  Adjustable filling tube horizontal 142-220 mm vertical 10-200 mm
  Telescopic filling tube for Automatic setting the length from 210 to 370 mm
Pipette
  Circulation valve for priming without loss of reagent
  Dispensing range from 0.01 to 10.00 ml

## Slide 14: PORTABLE WATER ANALYSIS KIT
Compact in a Briefcase, operational both on AC mains/Rechargeable battery 6V/12V for measurement of pH, EC, ORP, DO &Temperature, have LCD display & temp. compensation.
pH
  Range: 0-14 pH
  Resolution: 0.01pH
  Accuracy: + 0.01 PH
EC
  Range: 0.5 µS/cm to 100mS/cm (5 ranges)
  Resolution: 1 mV
  Accuracy: + 1%
Oxidation Reduction Potential (ORP)
  Range: 0-1999 mV
  Resolution: 1 mV
  Accuracy: + 1mV + 1 digit
DO
  Range: 0.1 to 20 ppm
  Resolution: 0.1 to 0.2 ppm
  Accuracy: + 0.2 ppm + 1 digit or 1% full scale deflection
Temperature
  Range: -5.0-99.9 0C
  Resolution: 0.10C
  Accuracy: + 1% 0C
Accessories: Electrodes for pH and EC measurements, cell holder, all standard buffers for 2 years, DO prove with membranes, temperature probe, one rechargeable compatible battery, instruction manual.

## Slide 15: FLAME PHOTOMETER
Flame photometer for routine application, microprocessor based with two filter for Sodium and Potassium ; Flame Atomization
4 Line 20 Character Alphanumeric LCD Display
Analysis of 1 to max. 4 elements in single aspiration
Automatic filter selection
Sensitivity: 2.0 ppm for Na, 1.0 ppm for K, 1.0 ppm for Li and 30.0 ppm for Ca
Facility for Re-standardization, Sample Data storage in memory and recall and Print out
Data transfer to Computer through RS-232 (PC-Link) (Optional)

## Slide 16: ION METER
Micro processor controlled multi-channel ion meter with LED display for menu, capable of measuring pH, milli volt, ORP (Oxidation Reduction Potential), Cations and anions, built in software
pH range: -2.000 to 19.999, Resolution: 0.001/0.01/0.1, Relative Accuracy: ± 0.005
Calibration point : 1-6 point Concentration range-0.000 to 19900.
Milli volt range : 1600 mV, milli volt accuracy: ±0.2 mV or ± 0.05 % of reading whichever is greater.
Working temperature: -50C to 105 0C and 5 to 85 % relative humidity, non-condensing.
Specific ion electrodes of Fluoride, Nitrate, Br etc. Fluoride electrode should have capability of measuring in conc. Range of saturated to 0.02 ppm at temp range 0 to 800C.
Installation of the instruments at any Regional laboratory of CGWB in India with AMC for 5 years after the warranty period.

## Slide 17: PC based UV-VIS SPECTROPHOTOMETER
Double beam, fully automatic UV-VIS Spectrophotometer and should also have a facility to connect with latest PC configuration for latest Windows XP base operation and compact with a fast laser printer
LCD display, 220 V AC, Stray Light better than 0.02% T at 220 nm (NaI) and 340nm (NaNO2)
Better baseline flatness less than + 0.0006 Abs. over the entire range of 190 to 1100nm.
Scan speed: Variable options of 3000nm/min to 2nm/min so as to capture maximum data points, hence to increase the resolution.
Noise level less than 0.00005 Abs. at 700nm.
Baseline stability less than 0.0003 Abs/Hour at 700nm
Multi-wavelength Photometric Measurement mode as standard
Abs / %T measurement possible at 6 wavelengths simultaneously in Photometric mode
Photometric, Multi-Wavelength Photometric, Spectrum, Quantitation, Kinetics and Multi-Component measurement modes should be provided.
Expanded Photometric range up to -4 to +4 Abs / 0 to 400%, enables high concentration measurement.
Photometric repeatability less than +/- 0.001 Abs at 1Abs.
Detector : Dual silicone photodiode
Monochromator having High end optics feature with Czerny-Turner mounting for high energy though put and high sensitivity
Spectrophotometer should have facility for 5 USB ports for data communication with PC and Printer compatibility.
Two sets of Quartz cell (optional)
USB Pen drive compatibility for data storage and transfer
Built-in- hardware validation for the parameters such as: Wavelength Accuracy, Wavelength Repeatability, Spectral Bandwidth, Baseline Flatness, Baseline Stability, Noise Level, Photometric Accuracy, Photometric repeatability & Stray Light
Semi-automatic testing: Interactive display for simplified testing of test items
Automatic Testing :Automatic measurement and pass/fail evaluation and printing of results
Installation of the instruments at Regional lab. of CGWB in India with AMC for 5 years after the warranty period.

## Slide 18: Water Double Distillation Apparatus
Quartz double distillation apparatus with automatic cut off system
Main boiler
A built-in heater enclosed in quartz glass tube
Housing auto water level adjuster
Bottom outlet for cleaning waste discharge which is fitted with Teflon screw cap valve
A coil condenser fitted on top for condensing vapors
Whole unit is mounted on epoxy coated metal box stand and covered with acrylic sheet cover for protection
Features:
  Temperature resistant
  Saving of electrical energy in every start
  Convenient bottom drain for easy drainage as accumulated solids impair the efficiency, cleaning is also possible
  Distilled water in cold condition
  Automatic start, cut-off & re-start in synchrony with water pressure/opening by electronic circuit

## Slide 19: TOTAL ORGANIC CARBON ANALYSER (TOC)
Compact PC controlled TOC analyzer for the determination of TOC, TIC, TC, NPOC and POC in water samples.
The system should comply with DIN, ASTM and USEPA standards.
The system should have high temperature combustion unit of up to 950 Deg C to ensure complete digestion of even stable compounds.
Sample feed should be through valve and septum free direct injection.
The injection volume should be variable from 50ul to 500ul.
The system should have high sensitive focused beam NDIR detector.
The measurement range should be from 50ppb to 30000 ppm without sample dilution.
The system should have facility to do multipoint calibration with single standard using different injection volumes.
The system should have provision to select automatically the optimum calibration curve for sample measurement.
The instrument should have self check system to control all parameters for device safety as well to ensure quality of results. The quoted software should be FDA 21 CFR PART 11 compliance. The system should be upgradeable for solid sample analysis with either integrated double furnace or with High temperature combustion module (1300 deg C temperature ).
The system should also have provision to upgrade for Total Nitrogen measurement using either electrochemical detector or chemiluminiscence detector.
The supplier should optionally offer Auto sampler with 60 positions and optional tray of more than 100 positions.
Complete set of spares and consumables necessary for operation should be offered.
Suitable latest PC and Printer needed for operation should be quoted by the supplier.

## Slide 20: FOUR CHANNEL WET CHEMISTRY ANALYSER
(Four parameters to be analyzed simultaneously)
The analyser should be modular with the following modules; Auto sampler, module, module holder, module with flow cell and filter, detector, interface and software.
The system should be up gradable into more simultaneous analysis possibility, and capable of adding more parameters in future.
The sample is to have continuous flow with air bubble Segmentation and the detection is colorimetric, Provision for other detectors like IR and UV and flame photometer should be available.
Auto sampler:
  Random access sampler; should have more than 100 sample cups; should have built in rinsing pump, for day-to-day easy maintenance.
  Sample cup capacity is to be up to 10ml
  Optional built-in dilutor.
  Separate cups for working standards.
Chemistry Unit :
  Should have capacity to hold upto 5 chemistry modules
  Should have two multi speed proportioning pumps, The two pumps should provide 32 tubings for reagent transportation to the analytical module.
  Should have air pump. Air pump should have built-in compressor.
  The multi speed pump should have speed adjustments for normal (analysis), fast (washing) and slow (stand by position, to save reagents).
Chemistry modules:
  Each parameter should have separate module and each module to have its own filter and flow cell. All components of analysis like dialyser, reactor etc (wherever applicable) should be provided within each module.
  Parameter for analysis: Chloride, Total Hardness, Calcium, Phosphate and Bicarbonate. Nitrate + Nitrite, Fluoride, Total Nitrogen,
  The analysis module should have its own flow cell and filter.
  Flow cell to be supplied with black coating to avoid interference of external light in the analysis.
  Flow cell is to be part of the chemistry module and should not be inside the photometer.
Detector: Two Dual channel Digital detector with high degree of accuracy and wide dynamic range; capable of measuring four parameters simultaneously.
The photometer detector head is to have an insert board and cable
  Digital detector with high degree of accuracy and wide dynamic range.
  Capable of measuring four parameters simultaneously.
  Wavelength selection by insert of interference filters.
  Wavelength range 340-1050 nm.
  Accommodation for 2 flow cells of 5 upto 50 mm path length each
  Software controlled automatic baseline and peak height settings.
  Software controlled lamp switching on / off
  32 bits high resolution absorption recording
  Detecting range up to 6.5 AU
  Signal / noise ratio <0.0003 AU
Software : The software is to be windows based and user friendly. Provision should be there for raw data storage. Automatic back ground correction should be possible;
PC & Printer : Suitable branded latest PC & colored laser printer should be offered.

## Slide 21: ATOMIC ABSORPTION SPECTROPHOTOMETER (AAS)
Computer operated double beam Atomic Absorption Spectrophotometer with high sensitivity, capability to analyze Flame Atomization with Air-Acetylene and Nitrous Oxide Acetylene Flame with Hydride Vapour Generator and future upgradeable to Graphite Furnace.
Wavelength range from 185 nm to 900 nm., Slit width 0.2 to 2.0 mm in steps.
High efficiency atomization achieved through new design Polypropylene nebulizer with Platinum-Iridium orifice and Ceramic impact bead for high sensitivity. High-end automation with fully motorized burner head up-down motion with automatic optimum height setting.
Deuterium and background correction modes (preferably Self-Reversal) should be available as standard.
Standard measurement modes include Flame continuous.
Auto selection of Hollow Cathode lamps with motorized 8 lamp turret.
Fully computerized gas flow control system with automatic setting of optimum fuel and oxidant flow rates.
Hallow Cathode lamp for each single elements to ensure maximum sensitivity, Cathode Lamps-Fe, Mn, Pb, Zn, Cu, Co, Cd, Cr, As, Ni, Hg, Se, Bi and Au.
Regulators- Nitrogen, Nitrous oxide and Air Acetylene.
Compressor- Oil Free.
VGA- Continuous flow base automatic vapour generation for hydride forming elements with peristaltic pump
Incorporates all safety measures as standard including the unique leak check on power 'ON', flame monitor, gas pressure monitor, inter-lock for wrong burner head use, Air - Acetylene / N2 O - Acetylene automatic change- over with flow rate monitor using optical sensor and drain tank level monitor to prevent flash-back.
Advanced 32-bit Windows 2000 based software for complete instruments control, data acquisition and data processing.
Installation of the instruments at any Regional laboratory of CGWB in India with AMC for 5 years after the warranty period.

## Slide 22: ION CHROMATOGRAPH (IC)
Electronics Type | Microprocessor controlled digital signal processing, autoranging
Cell Drive | 8 kHz square wave
Linearity | 1%
Resolution | 0.00238 nS/cm
Output Range | Digital Signal Range: 0-15,000 µS/cm, Analog Signal Range: 0-15,000 µS/cm
Noise, Wet | <0.2 nS at 23 µS/cm background <0.1 nS at 1 µS/cm background
Filter | Rise times 0 to 10 s, programmable
Sampling Rate | 1 to 100 Hz, user settable or automatic
Cell Temperature | 5 °C above DC upper zone temperature to 60 °C maximum. User settable, working range is identical to settable range.
Cell Temperature Stability | <0.001 °C
Cell Temperature Compensation | Default 1.7% per °C ; programmable from 0–3% per °C
Flow Cell Maximum Pressure | 10 MPa (1500 psi)
Flow Cell Volume | 0.7 µL (Analytical); 0.02 µL (capillary)
Cell Electrodes | Passivated 316 stainless steel. Compatible with MSA
Cell Body | Chemically inert polymeric material
Heat Exchanger | Inert, tortuous-path for low axial dispersion
Software | Chromatography Management Software, requires Microsoft Windows® 2000, Windows 7, 8.1 or 10
Automated Procedure Wizards | Yes, standard feature
System Smart Startup and Shutdown | Yes, standard feature
System Wellness and Predictive Performance | Yes, standard feature
Application Templates | Yes, standard feature
Multivendor Automation Support of 3rd Party | Yes, fully controls over 300 different instruments from more than 30 manufacturers, including GC, HPLC, and MS.
Customizable System Control Panels | Yes, standard feature
Signal Channels | Conductivity, Total Conductivity
Data Trending | Yes, all detector numerical parameters
System Status Virtual Channels | Yes, standard feature
Power Failure Protection | Yes, standard feature
System Trigger Commands and Conditionals | Yes, standard feature
Daily & Sample Audit Trail | Yes, standard feature
System Calibration Storage | Yes, factory, present, and previous. Completely user selectable

## Slide 23: INDUCTIVELY COUPLED PLASMA MASS SPECTROMETER (ICP-MS)
Mass spectrometry unit :
(1) Spectrometry unit Quadru-pole mass spectrometer
Hyperbolic rod made of molybdenum (main rod)
Cylindrical rod made of molybdenum (pre-rod)
Power supply frequency:2.5MHz
Mass number range 2 to 250
Resolution Max. 2 M is possible.
Mass number stability 0.1 a.m.u
(2) Detector Channel Thoron detector
Must be installed perpendicular to the quadrupole rod
(3) Counting method Analog/Pulse selection
Pulse measurement 24 bit counter
Max. Frequency 108cps
Analog measurement: Gain 2 step selection
  Integration time can be set for each element
Plasma ion source
(1) Torch unit Spray chamber
Plasma mini-torch
Coaxial type nebulizer
Peristaltic pump
Equipped with adjustment function of X, Y, Z, three axes torch position
CPU control for direction of torch axis
(2) RF power matching unit : Equipped with automatic matching circuit
(3) Gas controller 3-flowline method for plasma
Coolant gas 7.9 L/min
Plasma gas 1.5 L/min
Carrier gas 0.2 to 1.4 L/min
CPU controls carrier gas
RF power supply :
(1) Oscillator Quartz oscillator
(2) Frequency 27.120 MHz 0.05%
(3) RF output 0.2, 0.8, 1.0, 1.2 kW
(4) Output stability Within 0.3%
(5) RF circuit element Power MOS FET
(6) Control system Controlled by using CPU
(7) Lighting unit Full automatic ignition start
(8) PA unit cooling Water-cooling
(9) Equipped with safety function
Vacuum container and exhaust unit
(1) Vacuum container 3-stage differential exhaust system
(2) Exhaust system 1st stage Rotary pump
2nd stage Air-cooling type turbo molecular pump
3rd stage Air-cooling type turbo molecular pump

## Slide 24: INDUCTIVELY COUPLED PLASMA MASS SPECTROMETER (ICP-MS) continued…..
Data processing unit-hardware
CPU : Branded Core Duo with 2 GHZ or more
Memory capacity : 1 GB RAM
External memory : Hard disk 120GB : CD-R / W and USB
Display : 17-inch color Flat LCD display
OS : Microsoft Windows XP
LAN interface : 10BASE-T
Mouse : Optical Scroll mouse
Printer : Hp Laser Jet
Data processing unit-software
(1) Qualitative analysis Semi-quantitative determination value calculation
Isotopic pattern comparison function
(2) Quantitative analysis Calibration curve standard, standard addition method, single-point calibration curve method
Drift correction
Concomitant element correction
Internal calibration standard correction
Blank deletion
Weighing amount correction
Judgment of standards
(3) Isotopic ratio measurement Isotopic ratio calculation
(4) Report function
(5) File function Analysis data can be filed in the format that can be used on commercial spreadsheet software.
(6) Time axis measurement
(7) Wide range scan
(8) Start-up adjustment Lens voltage, mass axis, and resolution setting
Power supply
3-phase, 220/200V±10%, 60/50 Hz, 20A
Cooling water
2-system
･ Temperature ： 300C or less
･ Flow rate ： 2.0 L/min. or more
･ Outer diameter of faucet ： 15.5mm
Argon gas : minimum 3 cylinders with regulators should be supplied having
Purity of 99.95% or more
Important Information : Instrument should be supplied on working platform wherein all the other items which are not mentioned above like working standards, Fume hood, etc. should be quoted separately. Installation of the instruments and Training at Regional lab. of CGWB in India with
AMC for 5 years after the warranty period.

## Slide 25: Gas Chromatograph with Mass Spectrometer (GC-MS)
High performance Designed towards improved productivity, enhanced efficiency, and ease of operation
Column oven temp up to 450 0C
20-step Column oven temperature programming
Maximum column oven heating rate of 250 0C/min
High pressure sample injection possible with standard Split / Splitless injection port
Capability of installing 3 injection ports and 4 detectors
Instrument should be quoted with Dual packed column and one Capillary injector
Detector sampling frequency of 250 Hz (4 msec sampling time) helps accurate acquisition of sharp and narrow peaks and also enables fast GC analysis
Ultra High Speed GC capability – reduces analysis time with better resolution suitable even for narrow bore column of I.D. 0.1mm
Third Generation Advanced Flow Controller (AFC)
Carrier Head Pressure Control up to 970 kPa
Carrier Flow Control up to 1200 ml/min
Split Ratio Setting up to 9999.9 : 1
7 Step Pressure / Flow programming
All auxiliary gases like Air, H2, Make-up gas for FID, septum purge and digitally controlled through third generation Advanced Pressure Controller (APC) and Gas Cylinders.
Patented Constant Linear Velocity mode ensures maximum separation efficiency for any carrier gas. Hence even N2 gas can be used very efficiently for capillary column analysis
Large LCD window (240 x 320 dots) Graphic User Interface makes operation very easy and enables display of all GC set parameters, monitored values and also displays the acquired chromatogram
Built in GC System Check and Self Diagnostics function
Mass Spectometer: Quadruple (gold coated)/Ion trap Mass Spectrometer.
Cooling Unit: Thermostatically recirulating.
Dynode Detector: High energy dynode detector (preferably +/- 15 KV) for +/- Ion mass analysis
Unit Mass Resolution Ionisation: Unit mass resolution over entire range of 10-900 AMU. External ionization source having option for positive and negative ions with automatic switching from El to Cl.

## Slide 26: Gas Chromatograph with Mass Spectrometer (GC-MS) continued…..
Turbo Pump: Highly efficient pump (approx 200L/s- maintenance free, all accessories for safety controls and indicators.
Library for Pesticides: Latest library with spectra and 200 pesticides
Equipment Tuning: Both Auto and Manual.
Gas Control: Separate control for Helium gas flow.
Software Features:
  New 32/64-Bit, Windows NT Compatible GC Solutions software, Digital data acquisition from GC
  Full Qualitative and Quantitative processing capabilities, Multi Tasking, Multi Function capability for easy operation, User Management Functions , Full GLP / GMP support Multi level security with passwords, Audit Trail, System Check, Validation Assistant, System Suitability functions built-in QA / QC Functions with Pass / Fail judgements, All-in-one file configuration for easy method transfer, Customized Report generation, Networking capability
  Installation of the instruments and Training at Regional lab. of CGWB in India with AMC for 5 years after the warranty period.
FID :
  Temperature Range : up to 400 0C
  Minimum detected quantity : 3 pg/S (dodecane)
  Dynamic Range : 10 7
ECD :
  Temperature Range : up to 400 0C
  Minimum detected quantity : 0.1 pg/S (- BHC)
  Dynamic Range : 10 4
FPD :
  Temperature Range : up to 350 0C
  Minimum detected quantity :
  P : 0.5 pgP/S (Tributyl Phosphate)
  S :8 pgS/S (dodecane thiol)
  Dynamic Range : P :10 4 , S : 10 3

## Slide 27: TOC, Life Span & LOG BOOK
Terms & Condition:
  All the sophisticated instruments should be purchased only with demonstration from firm’s representatives at the site of their supply to a particular Institute etc.
  All instruments should be purchased only with AMC for minimum period of five years after the warrantee period.
  Chemists of CGWB should get operational training (in India or abroad) of the said instrument for a sufficient time so that they can handle the instrument independently.
Life Span of the Instrument:
  Instrument such as pH, EC meters, Flame photometer and Ion meter these instrument remain in good working condition for a period of 5-7 years with continuous use. However, the electrodes used with these instruments have a life span of two –three years and thereafter need replacement.
  For sophisticated instrument like AAS, GC-MS, IC P, Ion chromatograph, the life span is usually seven to ten years after which the instrument starts giving problem which do not get rectified even by the firm’s representative as per the past experience in some of the Laboratories where the instrument could not function satisfactorily even after spending huge amount on its new spares as advised by the authorized dealer.
Log Book: In order to observe day to day functioning and utilization of instruments it is to important to maintain Log Book register of the instrument where daily records of these instruments should be entered. This will give us idea about the running condition of the instrument and also the total volume of the work done by the instrument.